## Catalogue
DRAFT — UNVERIFIED (model knowledge + manufacturer sources + peer-review). Per-row 'Source Verification' = Verified / Draft / UNVERIFIED. NEVER use unverified data clinically without cross-checking IFU/SmPC/protocol document.
System/Kit | Manufacturer | Country/Region | Type | Anticoagulant | Centrifuge | Spin Protocol (g × min) | No. Spins | Platelet Concentration Factor | Leukocyte Profile | Key Growth Factors Retained | Activation Method | Final Volume (mL) | Fibrin Scaffold | Aesthetic Uses | Scalp & Hair Use | CE Mark | FDA Status | Spain/AEMPS Status | DEPA Classification | Evidence Class | Clinical Confidence | Key Cautions | Source Verification | Source URL | Source Type
INTRASPIN (L-PRF System) | Intra-Lock International / BioHorizons Implant Systems | USA (centrifuge German-made) | 2nd-gen platelet concentrate — Leukocyte- & Platelet-Rich Fibrin (L-PRF Choukroun); solid clot/membrane not liquid PRP | NONE (glass/silica-coated tubes; no anticoagulant) | IntraSpin fixed-angle centrifuge | ~400–408g (2700 rpm) × 12 min | 1 | Not expressed as ×; platelets+leukocytes trapped in fibrin matrix | Leukocyte-RICH (LR) | PDGF, TGF-β1, VEGF, IGF + sustained release from fibrin | None exogenous — glass-contact clotting (no thrombin/CaCl₂) | L-PRF clots/membranes (solid; not a mL volume) | Yes | Off-label facial 'PRF' biostimulation, sub-dermal, sticky-bone in facial surgery | Yes — PRF microneedling/injection for hair | Yes (CE-marked) | FDA-CLEARED (only FDA-cleared device for PRF production) | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — técnica cerrada SOLO si administración directa del coágulo; operaciones adicionales = sistema abierto; ref ISS220 | Not applicable (DEPA is for liquid PRP not PRF) | Class II–III (large dental/OMF literature) | Moderate–High (oral/maxillofacial); Low for aesthetic/hair | Closed-technique status is CONDITIONAL — any manipulation beyond direct clot use reclassifies as open system | Verified | https://www.biohorizons.com/products/IntraSpin | Manufacturer + consensus guideline
3E-PRP (3E-10 / 3E-20) | Pervice Korea Co. Ltd. | Rep. de Corea | All-in-one syringe-type closed kit; single spin; buffy-coat (push-up recovery) | Not in kit (operator adds ACD-A/citrate) — type not confirmed | Standard swing-bucket (not supplied) | Not confirmed (g/min not published; ~10 min total process) | 1 | 5–10× baseline (mfr claim) | Likely LR (buffy-coat); not independently confirmed | PDGF, TGF-β1, VEGF, EGF, bFGF, IGF-1 (generic) | Optional CaCl₂/thrombin (operator) — not confirmed | Not confirmed (draw ≈10/20 mL per model; 3E-10/3E-20 = blood volume) | No | Facial rejuvenation / skin regeneration | Yes — AGA/hair-loss (mfr-listed) | Yes — CE 2015 (mfr timeline); KGMP | Mfr claims US FDA 2016 — NOT verified in 510(k) DB | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (iniciales); ref 3E-10; 3E-20 | Not classifiable (insufficient published data) | Class IV (regional brand; no independent RCT) | Low | Anticoagulant + loading needle NOT in kit → protocol variability; leukocyte profile unstandardized; FDA claim unverified | Verified | http://www.pervice.net/en/company/history | Manufacturer site + PR
Orthokine (Autologous Conditioned Serum — NOT PRP) | Rhein Device GmbH (tech: Orthogen AG) | Germany | Autologous conditioned serum (ACS): cell-free serum from whole blood incubated in glass-bead syringe; anti-inflammatory cytokines | None (whole blood incubated; no anticoagulant) | EOT II syringe / Orthogen device (CrSO₄-coated glass beads) + centrifuge for serum separation | Incubate 37°C 6–9 h (up to 24 h) then single separation spin | 1 (post-incubation) | N/A — not a platelet product | N/A (cell-free serum) | IL-1Ra (primary), IL-4, IL-10, IL-13, TGF-β1 (anti-inflammatory profile) | Monocyte stimulation by glass beads during 37°C incubation | Several mL serum aliquots from 50–60 mL draw | No | Off-label; not a standard aesthetic PRP | Off-label only (not established) | Yes (CE-marked EU device) | FDA IDE approved Jan 2023 for a knee-OA trial — NOT FDA-cleared for commercial use | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (finales); ref F0005 | Not applicable (not PRP) | Class II (RCTs in knee OA/radiculopathy) | Moderate–High (OA); Low/None for aesthetics-hair | MISCLASSIFICATION RISK: this is ACS not PRP — different mechanism and indications | Verified | https://pubmed.ncbi.nlm.nih.gov/20441419/ | Peer-review reviews + RCT
WINT PRP (DKK-1/Wnt joint product — NOT aesthetic PRP) | Biovico Sp. z o.o. | Poland (Gdynia) | Autologous DKK-1-enriched blood derivative — borosilicate activator triggers Dickkopf-1 release to inhibit Wnt (for intra-articular OA) | Not in kit | Dedicated centrifuge (Xerthra-type platform) | Not confirmed | Not confirmed (single typical) | Not confirmed | Not confirmed | DKK-1 protein (target) + platelet growth factors | In-tube borosilicate activator | Not confirmed (intra-articular use) | No | NOT indicated for aesthetics (joint/OA product) | No | Yes (CE-marked) | Not confirmed | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima; ref PRO00579 | Not classified | Class IV | Low | Different purpose from Xerthra: WINT = DKK-1/Wnt-inhibition joint product not an aesthetic/hair kit | Draft | https://biovico.com/product | Manufacturer/distributor listing
Autologous Exosome / EV Kits (class) | Multiple manufacturers — no single dominant brand verified for clinic-ready autologous exosome preparation | Global (mostly Asian + European startups) | Autologous EVs (claimed exosomes — see Cautions) | Standard ACD-A or citrate (blood draw) | Ultra-high-speed centrifuge (100,000g) required for true exosomes — NOT achievable with clinic centrifuges. Commercial kits use precipitation reagents (e.g. polymer precipitation) or size-exclusion columns instead. | True exosome isolation: 300g × 10 min + 2000g × 10 min + 10,000g × 30 min + 100,000g × 70 min (NOT achievable in clinic). Commercial kit approximations: 3000g × 15 min + precipitation step (kit-dependent) | 3–4 (minimum) | N/A (exosome concentration measured in particles/mL, not platelet factor) | N/A (exosomes are cell-derived vesicles 30–150 nm; do not contain leukocytes per se) | miRNAs, tetraspanins (CD9, CD63, CD81), growth factors encapsulated in vesicles — highly dependent on cell source and protocol | None needed for EVs | Extremely variable — protocol-dependent (typically < 0.5 mL) | No | Proposed: skin rejuvenation, hair, wound healing — largely preclinical or early case series | Very limited published data | CE status: no approved autologous exosome kit identified | FDA: No cleared autologous exosome kit for aesthetic use identified | Spain/AEMPS: No specific approval identified | Not classified under DEPA (EVs, not PRF/PRP) | Class IV / Preclinical (minimal clinical evidence in aesthetics) | Very Low — theoretical interest; commercially available kits producing 'true' exosomes at clinic level are not validated; most products are EV-enriched fractions, not pure exosome preparations | CRITICAL: (1) Most commercial 'autologous exosome kits' produce EVs/nanoparticles, not International Society for Extracellular Vesicles (ISEV)-compliant exosomes. (2) True exosome isolation requires ultracentrifugation not feasible in clinics. (3) Autologous exosomes are distinct from allogeneic (e.g. ExoCoBio ASCE+, Purasomes — Catalog 4). (4) Regulatory gray area — no country has cleared a clinic-ready autologous exosome kit. (5) Evidence in aesthetics: preclinical only or very early phase. | UNVERIFIED | ISEV Guidelines 2018 (Théry et al. J Extracell Vesicles 2018) for definition standards | Peer-review (for EV definition); commercial products = UNVERIFIED
TriCell Best (TriCell New / New S) | REV-MED Inc. | Rep. de Corea | Autologous PRP — 3-chamber double-spin buffy-coat closed system | ACD-A (3 mL ACD-A / 30 mL blood per study) | Standard bench centrifuge | Not confirmed (per IFU) | 2 (double) | ~3.02× (peer-reviewed ophthalmic protocol) | Selectable PRC (LR) / PLC (leuko-lite) / PCP (conditioned plasma) | PDGF, TGF-β, VEGF, EGF, IGF (generic) | Optional (buffy-coat) — not confirmed | 1–4 mL (typically 2–4 mL) | No | Aesthetic/cosmetic (brand) | Yes (aesthetic/scalp marketed) | Yes — CE + ISO 13485 (also KFDA/NMPA/PMDA) | Not confirmed (no US 510(k) found) | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima; ref Tricell New TEQAGRU0001 / New S TEQAGRU0002 | L-PRP or P-PRP (adjustable) | Class III–IV (lab/animal + Korean MOHW new-tech 2024 for knee OA) | Low–moderate | Kit does NOT include PRP-extraction syringe (must be supplied to keep closed); ~3× is an ophthalmic protocol not aesthetic | Draft | https://rev-med.com/portfolio/tricell-prp/ | Manufacturer + peer-review
PRO PRP (P10 / P30A / P60) | Goodmorning Bio Co. Ltd. | Rep. de Corea | Autologous PRP — buffy-coat single-spin closed kit | NOT in kit (clinic supplies; typically ACD-A/citrate) | Standard bench centrifuge | Not confirmed (single spin — avoids double-spin platelet damage) | 1 | 12–13× (vendor claim; unverified) | LR (buffy-coat collected to syringe) | PDGF, TGF-β, VEGF, EGF, IGF (generic) | Proprietary mechanical (physical) activator — no chemical | Not confirmed (kit sizes 10/15/20/30/60 cc) | No | Skin rejuvenation (brand) | Yes | Yes — 'CE recognized', gamma-sterilized; exports 30+ countries | Company states FDA 510(k) cleared (CBER; specific number not confirmed) | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima; ref P30A; P10; P60 | L-PRP (buffy-coat; LR) | Class IV (regional brand; no human aesthetic RCT) | Low–moderate | Kit excludes anticoagulant (must be supplied to keep closed classification); '12–13×' unverified | Draft | http://goodmorningbio.com/eng/ | Manufacturer
PROLO 30 PRP | RM Life Science Co. Ltd. | Rep. de Corea | Autologous PRP — closed single-syringe kit (kit completo) | Citrate-based (included) | Standard bench centrifuge (kit incl. counter-balance) | Not confirmed | Not confirmed (likely single) | 4–6× baseline (brand claim) | Not confirmed (buffy-coat → likely LR) | PDGF, TGF-β, VEGF, EGF, IGF (generic) | Optional CaCl₂ — not confirmed | Not confirmed (30 cc syringe kit) | No | Wrinkles, periorbital dark circles, acne marks, loose skin (brand) | Yes — hair loss listed | Not confirmed (RM at MEDICA 2024/25; 'CE' claimed on trade listings; cert not verified) | Not confirmed | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima; ref A001-0101 | L-PRP (likely; unverified) | Class IV (regional brand; no independent RCT) | Low | Multiple Korean firms sell 'Prolo' kits (RM; N-Finders; Tracle) — do not conflate; specs from trade listings not an RM IFU | Draft | https://www.ec21.com/product-details/PRP-Kit-PROLO-30--9732882.html | Manufacturer/trade listing
RegenKit-THT | Regen Lab SA | Le Mont-sur-Lausanne, Switzerland | Autologous Thrombin | Citrate | Standard bench centrifuge | 900g × 5 min | 1 | N/A — thrombin activator kit | N/A | Thrombin (endogenous activator) | Provides autologous thrombin for activating other Regen Lab kits | ~0.5 mL thrombin (from 3 mL blood) | Activates BCT/BCT-3 into gel | Used as activator only — not standalone treatment | N/A | Yes — CE | Registered | Available in Spain | Activator kit only | N/A | N/A (activator) | Do not use as standalone PRP injection; thrombin concentration varies; must be used immediately post-preparation | Draft | https://www.regenlab.com/products/ | Manufacturer
VIVOSTAT PRF | Vivostat A/S | Denmark | Automated autologous PRF + fibrin sealant (sprayable) — 1st-gen PRF | Citrate (supplied; added in Preparation Unit) | Vivostat System — automated Processor Unit + Applicator/Spraypen (not a manual centrifuge) | Automated proprietary process (~23–30 min) | Automated (single integrated process) | ~7× baseline (>1,000,000/µL) | Not confirmed (plasma-based; likely leukocyte-reduced) | TGF-β1, PDGF, VEGF, FGF-2 | No separate thrombin — fibrin polymerizes on spray application | 5–6 mL PRF from 120 mL blood | Yes (sprayable fibrin matrix) | Limited/off-label aesthetic (surgical wound/tissue regeneration is core) | Not confirmed | Yes (CE-marked EU device) | Not confirmed (no clear US PRF clearance found) | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (iniciales y finales); ref VS406; VS400; VS410; VS422 | Not classified | Class III (surgical/wound literature) | Moderate (surgery/wounds); Low for aesthetics | High blood draw (120 mL); positioned for surgery/haemostasis not office aesthetics | Verified | https://vivostat.com/vivostat-prf/ | Manufacturer + peer-review pilot
CGF — Medifuge | Silfradent s.r.l. | Santa Sofia (Forlì-Cesena), Italy | CGF (Concentrated Growth Factors) | None — glass tubes (auto-activation on blood-glass contact) | Medifuge MF-200 / MF-200plus (proprietary variable-speed centrifuge — mandatory; CGF protocol REQUIRES this centrifuge) | Variable speed program: 30 s @2700 rpm → 2 min @2400 rpm → 4 min @2700 rpm → 3 min @3000 rpm → 3 min @2400 rpm (deceleration) — total ~13 min | 1 (variable speed) | Very high in fibrin matrix; serum separated for injectable use; CD34+ stem cell enrichment claimed | LR — very high leukocyte and CD34+ stem cell content (key differentiator from standard PRP/PRF) | TGF-β1, IGF-1, VEGF, BMP-2/7 (variable), CD34+ cell-derived factors — claimed to exceed PRF by density | Auto-polymerization (no exogenous activator) | CGF membrane (fibrin block) + ~2–4 mL injectable serum (per 10 mL blood) | Yes — very dense fibrin membrane (denser than L-PRF) | Wound healing (post-implant), facial volume (membrane + serum injection), hair transplant (scalp, donor site), injectable serum for rejuvenation, combination with HA/filler, dental (primary indication) | Yes — hair transplant surgery (donor/recipient wound); some AGA serum protocols | Yes — Silfradent CE marked (EU) | Not FDA-cleared for clinical use in USA aesthetic | Available in Spain (CE device; aesthetic use off-label) | Not formally classified under DEPA; closest: L-PRF variant with very high fibrin density | Class III (mostly Italian case series and small cohorts; no multi-centre RCTs vs PRP) | Low-moderate (strong Italian/Mediterranean use; CD34+ cell claim is interesting but not consistently validated; requires Medifuge centrifuge — protocol cannot be reproduced on standard centrifuge) | Requires Silfradent Medifuge centrifuge (not substitutable — variable speed profile is proprietary); USA: not FDA-cleared; evidence base primarily Italian with limited international replication; CD34+ concentration claim requires independent validation; injectable serum composition varies | Draft | https://www.silfradent.com/medifuge/ | Manufacturer
QREM CYTOKINE (autologous cytokine serum — NOT PRP) | Tecnología Regenerativa QREM SL | Spain (Barcelona/Catalonia) | Closed automated two-chamber Q-Kit + Q-Device ('lab-in-a-box') → cell-free cytokine-rich serum (SARC) | Sodium citrate (~2 mL chamber F) — AEMPS: kit no incluye anticoagulante ni jeringas/conectores | Integrated in Q-Device (automated) | Automated 8-stage cycle ~30–36 min | Multi-step automated | N/A (final product has no platelets) | N/A (cell-free serum) | Anti-inflammatory + anabolic cytokines/GFs (analytes not published) | Biological (endogenous activation → transient fibrin; no exogenous CaCl₂) | Not confirmed (from ~40 mL blood) | No | Dermatology (scar correction, chronic wounds); primary line musculoskeletal (OA, tendinopathy) | Not confirmed | Yes — CE MDR00040 (KIWA 0476) Class IIa; AEMPS mfg licence 7096-PS | Not confirmed / No | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (iniciales); ref M01-001; M02-001 (NOTA: es un suero de citoquinas, no un concentrado de plaquetas) | N/A (not PRP) | Class III–IV | Low | DO NOT treat as PRP — no platelet concentrate delivered; final volume/dosing unpublished; aesthetic/hair indication limited | Verified | https://qrem-regenerative.com/preguntas-frecuentes/ | Manufacturer + KIWA CE + AEMPS
OLIN-1 (20/40 mL) | Enelini Associats, SLU | Spain (Barcelona) | Closed single-device system (draw + spin + fraction separation in one sterile device; separates PPP) | Not in kit (distributor-supplied) | Not confirmed (no specific centrifuge named) | Not confirmed | Not confirmed (likely 1) | Not confirmed | Not confirmed | PDGF, TGF-β, VEGF, EGF, IGF-1 (generic — device data not confirmed) | Optional CaCl₂ (standard practice) — not confirmed | Not confirmed (PRP + PPP from 20/40 mL) | No (liquid unless activated) | PRP regenerative/aesthetic (general) | Not confirmed | Yes (CE + ISO; AEMPS mfg licence) | Not confirmed / not FDA-cleared (EU device) | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — técnica cerrada (si jeringa anticoagulante precargada + conector) / exposición mínima (iniciales); ref OL-1999; OL-1994 | Not confirmed | Class III–IV | Low | Concentration factor / spin / leukocyte profile not published — verify via IFU | Draft | https://enelini.com/en/olin-1-2 | Manufacturer + AEMPS registry
SW-PRP (Kit cerrado para PRP) | Seawon Meditech Co. Ltd. | Rep. de Corea | Closed single-tubing system (container + hose; two extraction pathways); buffy-coat | Not in kit (operator adds) — type not confirmed | Standard swing-bucket (dedicated unit exists; not in kit) | Not confirmed | Not confirmed (single presumed) | Not confirmed | Likely LR (buffy-coat); not confirmed | PDGF, TGF-β1, VEGF, EGF, bFGF, IGF-1 (generic) | Optional CaCl₂/thrombin (operator) — not confirmed | Not confirmed (~20 mL capacity) | No | Facial/skin rejuvenation + orthopedic | Plausible — not confirmed | Yes — CE 2013 (mfr); KFDA 2012 | Not confirmed (KFDA + CE only) | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (iniciales); ref SW-PRP | Not classifiable | Class IV (regional brand; no independent RCT) | Low | Anticoagulant + needles + syringes NOT in kit; spin/concentration/leukocyte data unpublished; no FDA | Verified | http://www.seawonmt.com/en/products/aesthetic/sw-prp/ | Manufacturer site + catalogue
Hy-Tissue PRP (PRP 20 / PRP 50) | Lab. Fidia Farmacéutica SLU (parent Fidia Farmaceutici SpA) | Spain / Italy | Closed single-use kit; nano-polished anti-electrostatic tube + laminar-flow 'Push.out' selection | Not in kit; distributor supplies C-Lock 4% sodium citrate 5 mL precharged syringe | Not confirmed (soft-spin; model not published) | ~8 min single soft spin (exact g not confirmed) | 1 | Up to ~2.9× baseline; ~97% platelet recovery | Two selectable: very-pure (LP) or leuco-rich (LR) | PDGF, TGF-β, VEGF, EGF, IGF-1 (generic) | None default (laminar-flow; optional CaCl₂) | 4 mL (PRP 20) / 10 mL (PRP 50) | No | Skin rejuvenation, hyperpigmentation, scars/fibrosis, wound healing, genital atrophy | Yes — MAGA/FAGA & hair-transplant graft survival | Yes (CE) | Not confirmed for Hy-Tissue (Fidia's FDA-listed PRP is Richplat Ortho) | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — técnica cerrada (si jeringa C-Lock 4% + conector) / exposición mínima (iniciales); ref P7-4050; P7-4020 | Not confirmed (high purity; no activation; moderate dose) | Class III–IV (no device-specific RCT confirmed) | Low | Exact g not published; modest ~2.9× vs buffy-coat systems; confirm LP vs LR per indication | Verified | https://www.hy-tissue.com/es/aes-hy-tissue-prp/ | Manufacturer + AEMPS registry
BS PRP System (BS PRP-20) | BS Medical Co. Ltd. | Rep. de Corea | Cylinder + screw-rod closed separation device (not gel tube); buffy-coat | Not in kit (operator adds) — type not confirmed | Standard swing-bucket (not supplied) | Not confirmed | Not confirmed (single presumed) | Not confirmed ('higher concentration' claimed) | Likely LR (buffy-coat); not confirmed | PDGF, TGF-β1, VEGF, EGF, bFGF, IGF-1 (generic) | Optional CaCl₂/thrombin (operator) — not confirmed | Not confirmed (~20 cc draw) | No | Aesthetic/skin + orthopedic | Not explicitly listed — not confirmed | CE implied (AEMPS-listed; exports ~70 countries) — not confirmed | Not confirmed | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (iniciales); ref BS PRP-20 | Not classifiable | Class IV (regional brand; no independent RCT) | Low | Anticoagulant + extraction system NOT in kit; spin/concentration unpublished; CE/FDA unconfirmed | Draft | https://bsmedical.gobizkorea.com/mini/site/miniSiteMain.do?domn_id=bsmedical | Manufacturer minisite + marketplace
i-PRF (Injectable PRF) | Protocol: Miron RJ / Choukroun J — not a commercial brand (requires compatible centrifuge + plastic tubes) | Protocol origin: France/Germany | i-PRF | None — plastic/silicone tubes used (absence of glass prevents immediate platelet activation on draw) | Any centrifuge capable of very low-speed: 60g × 3 min | 60g × 3 min (ultra-low speed protocol — Miron 2017) | 1 | ~2–3× estimated (difficult to measure pre-gelation; lower than standard PRP because very slow spin) | Low-moderate LR — buffy coat partially included at this low spin; less leukocyte activation vs high-spin LR | PDGF, TGF-β1, VEGF, FGF (fibrin-bound — higher growth factor entrapment per mg fibrin vs PRP) | None exogenous — fibrinogen begins polymerizing ~10–15 min post-centrifuge at body temperature; inject immediately after centrifuge | 1–2 mL (upper yellow layer from 10 mL blood — small volume) | Yes — forms fibrin scaffold in situ after injection (this is the key feature) | Mesotherapy-style skin rejuvenation, periorbital area, under-eye, skin texture, hair (AGA — off-label), combination with PRP or HA filler | Yes — growing AGA use (horizontal centrifugation variant); combination with PRP for hair | No specific CE product (protocol-based — requires CE centrifuge + CE tubes) | No specific FDA product (protocol-based) | Spain: widely used clinically; protocol-based (no single regulated product) | P-PRF (fibrin forming) or ambiguous: between P-PRP and P-PRF | Class III (mostly case reports and small series; emerging RCTs 2022–2025) | Low (novel protocol with promising mechanism; insufficient RCT evidence; technique-sensitive; polymerization timing critical) | Very small volume; timing from centrifuge to injection critical (< 15 min or begins to gel in syringe); speed must be very precise (not all centrifuges achieve 60g reliably); protocol lacks standardisation across practices; limited evidence vs standard PRP | Draft | Miron RJ et al. 2019 (J Clin Periodontol); Choukroun J + Miron RJ — multiple publications | Peer-review / Protocol
Prosys ACT10 Bio Kit | Medisarang Co. Ltd. (AEMPS legal mfr); brand PROSYS by Prodizen Inc. — same-entity link NOT confirmed | Rep. de Corea | Low-volume closed PRP tube (PRP from 10 cc blood); single spin; buffy-coat (no gel) | Included (kit completo) — exact agent (ACD-A vs citrate) not confirmed | Any standard centrifuge (Prodizen TD5/TD4C separate) | 4000 rpm × 5 min (g not stated) | 1 | 6–8× baseline (PROSYS line; ~2 mL PRP) — ACT10-specific not confirmed | Likely LR (buffy-coat); not confirmed | PDGF, TGF-β1, VEGF, EGF, bFGF, IGF-1 (generic) | Optional CaCl₂/thrombin (operator) — not confirmed | ~2 mL PRP from ~10 cc blood (ACT10 = 10 cc); gamma-irradiated | No | Skin regeneration / facial rejuvenation | Yes — hair/scalp (PROSYS line) | CE for PROSYS line (also KFDA 2008/Japan/Taiwan/ISO); ACT10-specific cert not confirmed | ACT10 not confirmed (Prodizen 510(k) is for a different product 'WORLD PRP') | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (iniciales y finales); ref Prosys ACT10 | Not classifiable | Class IV (regional brand; no independent RCT) | Low | Manufacturer discrepancy Medisarang vs Prodizen — verify; ACT10-specific concentration/CE/FDA not individually documented | Verified | https://prosysglobal.com/en/ | Manufacturer site + distributor
DISPRAS | Lab. Fidia Farmacéutica SLU | Spain / Italy | Not confirmed — Fidia P7-series closed PRP blood-separation kit (same P7 prefix as Hy-Tissue) | Not in kit (distributor-supplied) | Not confirmed | Not confirmed | Not confirmed | Not confirmed | Not confirmed | Not confirmed | Not confirmed | Not confirmed | No | PRP use (general; not confirmed) | Not confirmed | Yes (CE; Fidia device) | Not confirmed | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (iniciales); ref P7-1212; P7-1222; P7-1232; P7-1242; P7-1212-SB; P7-1222-SB | Not confirmed | Class III–IV | Low | No public specs — obtain IFU from Fidia; '-SB' suffix meaning unknown (possible skin-booster variant) | UNVERIFIED | https://www.aemps.gob.es/medicamentos-de-uso-humano/medicamentos-biologicos/plasma-rico-en-plaquetas-prp/listado-de-productos-sanitarios-considerados-como-sistemas-de-obtencion-de-prp-cerrados-comunicados-en-la-aplicacion-ccps/ | AEMPS registry only
Plasmagel (Autologous Thermally-Gelled Plasma) | No commercial kit — operator-assembled protocol (autologous procedure using standard blood collection + heating device) | Protocol origins: Redaelli A (Italy); various European practitioners | Plasmagel (Autologous filler — thermal denaturation of plasma proteins) | ACD-A or sodium citrate (standard vacutainer draw) | Standard centrifuge for plasma separation; thermal block / water bath for gelation step | Plasma separation: 1500–2000g × 8–10 min. Gelation: heat isolated plasma to 74–76°C × 20–30 min in sterile closed container | 1 (centrifugation) + thermal step | N/A — platelets destroyed/denatured by heat; product is fibrinogen-based gel | N/A — leukocytes denatured | N/A — all growth factors denatured by thermal treatment. Mechanism is PURELY VOLUMETRIC (not regenerative). Key proteins: denatured fibrinogen, fibronectin, albumin forming gel matrix. | Thermal denaturation at 74–76°C × 20–30 min | 3–5 mL gel (from 10 mL blood → 5–6 mL plasma → 3–5 mL gel after water loss) | Yes — the product IS a semi-solid thermal fibrin gel | Autologous filler for: lips, nasolabial folds, tear trough, volume restoration (small zones), skin booster layer | No scalp/hair use (not a growth factor product — heat destroys GFs) | Not applicable (autologous procedure; no commercial product to CE-mark; heating device used must be CE-marked) | Not applicable (autologous procedure — classified as 'same surgical procedure' in EU/US; not regulated as a drug/device when used autologously in same session) | Spain: Used by some practitioners; regulatory status = autologous procedure (not product); falls under surgeon responsibility | Not classified under DEPA (not a PRP/PRF product — thermal gel) | Class III (very limited: case series, retrospective cohorts only) | Low (interesting autologous concept; results unpredictable; longevity 3–6 months but highly variable; no large RCTs; falling out of favour vs HA + PRP combined approaches) | CRITICAL DISTINCTION: Plasmagel is NOT regenerative — all growth factors are destroyed by heat. It is a volumetric autologous filler only. Longevity shorter and more variable than HA. Thermal protocol must be precise (< 74°C → insufficient gelation; > 78°C → protein breakdown + discolouration). Patient must understand autologous limitations. Sterility of heating step critical. Minimal modern clinical data. | Draft | Redaelli A et al. — case series (Aesthetic Plast Surg); Valli R (Aesthetic Med) | Peer-review (limited)
L-PRF / A-PRF (Leukocyte- / Advanced-PRF) | Protocol origin: Choukroun J (L-PRF 2001); Miron RJ (A-PRF 2014). Commercial devices: Intralock PRGF-Endoret (separate from BTI), Process PRF (Nice), IntraSpin (Intralock/BioHorizons) | France / USA | PRF (L-PRF / A-PRF) | None — glass tubes; blood-glass contact immediately activates fibrin polymerization on draw | Specific validated centrifuges: Process PRF by Process (Nice, France); IntraSpin by Intralock (Boca Raton, FL) | L-PRF: 2700 rpm (~408g) × 12 min ¦ A-PRF: 1300–1500 rpm (~200g) × 14 min | 1 each | Embedded in fibrin membrane — not measurable as liquid concentrate; platelet recovery ~97% within membrane | LR membrane (rich in leukocytes, CD34+ stem cells, CD66b+ neutrophils) | TGF-β1 (very high in membrane), PDGF-AB, VEGF, IGF-1, FGF-2 — released over 7–14 days (sustained release) | Auto-polymerization (no exogenous activator — triggered by glass contact) | L-PRF: ~1 fibrin membrane per tube (10 mL blood/tube) — not injectable ¦ Liquid PRF serum collectable ~0.5 mL/tube | Yes — dense fibrin membrane (L-PRF/A-PRF are membranes, not liquids) | Wound healing, post-procedure recovery, hair transplant surgery (scalp coverage), sutured into incisions, PRF exudate for mesotherapy (membrane squeezed = liquid PRF serum) | Yes — hair transplant wound coverage; PRF serum injected peri-follicular; some AGA protocols combine PRF serum with PRP | CE marked (Process PRF, IntraSpin) | IntraSpin FDA 510(k) cleared (K161166 — bone graft/dental) | Available in Spain (primarily dental/oral surgery; aesthetic use off-label) | L-PRF: Leukocyte-rich, High fibrin density | Class II (dental/implant); Class III (aesthetic) | Moderate for wound/surgical; Low for injectable aesthetic (product is membrane, not liquid — fundamentally different application vs PRP) | NOT injectable as standard preparation — liquid serum yield very small; requires validated centrifuge (not all bench centrifuges compatible); primarily dental/surgical indication; aesthetic injection = membrane-squeezed serum (variable composition); A-PRF claims more growth factors than L-PRF but evidence is limited | Verified | Choukroun J et al. Platelet-Rich Fibrin (PRF): A Second-Generation Platelet Concentrate. Part IV: Clinical Effects (Oral Surg 2006); Miron RJ et al. A-PRF JCDA 2014 | Peer-review / Protocol
Selphyl PRFM (Platelet Rich Fibrin Matrix) | Regenity Biosciences (formerly Aesthetic Factors LLC) | Parsippany, NJ, USA | PRFM (Platelet Rich Fibrin Matrix) | None in collection tube (citrate-free draw); CaCl₂ added post-centrifuge to initiate fibrin polymerization | Standard centrifuge | 1100g × 6 min | 1 | ~1.5–2× baseline (fibrin matrix traps platelets — concentration in matrix higher than liquid) | LP (very low leukocyte content) | PDGF-AB, TGF-β1, EGF, VEGF (trapped in fibrin scaffold — sustained release claimed) | CaCl₂ added post-centrifuge → autonomous fibrin gelation within 8–12 min | 0.5–1 mL fibrin matrix (from 3 mL blood — volume limited) | Yes — product IS a fibrin matrix (PRFM = platelet-rich fibrin matrix); cannot be injected after full gelation — must be injected during gelation window | Facial rejuvenation (NLF, crow's feet, lips, tear trough, cheeks — small volume zones), post-filler/toxin, skin texture | Limited scalp/hair data (volume too small for full scalp coverage) | Yes — CE | FDA 510(k) cleared (K060477 — blood processing) | Available in Spain via distributors | P-PRF: Pure, High fibrin density | Class III (several case series; few small RCTs with mixed results) | Low-moderate (unique fibrin matrix format; small volume limits use; sustained release claim not consistently validated; results variable vs standard PRP) | Very small volume limits indications; inject during gelation window (timing critical); fibrin immediately dilutes in tissue; limited to small zone treatments; more expensive per mL than standard PRP | Draft | https://regenity.com/selphyl/ | Manufacturer
PRGF-Endoret | BTI Biotechnology Institute (Anitua EA) | Vitoria-Gasteiz, Spain | PRGF | Sodium citrate 3.8% | Standard bench centrifuge | 580g × 8 min | 1 | 1.4–1.7× baseline (modest — by design, avoids over-concentrating pro-inflammatory fractions) | LP (Leukocyte-Poor — intentional design philosophy: excludes leukocytes to minimise inflammatory cytokines) | VEGF, EGF, PDGF-AB, IGF-1, FGF (all preserved in plasma fraction) | CaCl₂ (50 µL per 1 mL PRGF) + 37°C incubation → liquid PRGF or fibrin membrane depending on processing | 1–2 mL injectable (plasma fraction only, from 9 mL whole blood) | Yes — CaCl₂ activation at 37°C produces fibrin membrane for topical/sutured use | Facial rejuvenation (skin quality, fine lines), periorbital, tear trough, neck, décolleté, post-laser/peel healing, ophthalmic (dry eye — strongest evidence), dental/implant | Yes (hair loss, aesthetic off-label; some published series) | Yes — CE-marked in EU for autologous use | FDA 510(k) cleared as blood processing device; specific aesthetic indications not FDA-cleared | Widely used in Spain; BTI is Spanish origin; CE-marked — compliant for ES use ¦ AEMPS CCPS 17-feb-2025: kits ENDORET (KMU10 y KMU10-TPC estéticos; KMU11/KMU15/KMU18/EDK1/KMU19 no estéticos) — exposición mínima (finales); ver hoja 09 | P-PRP (liquid form) / P-PRF (fibrin membrane form): Pure/LP, Low fibrin density | Class IIb (orthopedic/ophthalmic); Class III for aesthetic (no large RCTs) | Moderate for ophthalmic; Low-moderate for aesthetic (compared to other PRP systems the LP design has less platelet yield but arguably cleaner GF profile) | Platelet yield lowest of all systems by design; activation timing critical; fibrin membrane not injectable; distinct from PRP — do not substitute protocols; no generics | Draft | https://www.bti-implant.es/endoret/ | Manufacturer
KIT ENDORET ESTÉTICA (BTI PRGF-Endoret) | BTI Biotechnology Institute SL (E. Anitua) | Vitoria-Gasteiz, Spain | PRGF (P-PRP; leukocyte-poor) — application-specific kit | Sodium citrate 3.8% (0.4 mL/9 mL tube) | BTI PRGF System V (+ Plasmaterm H for gel) | ~580g × 8 min (RT) | 1 | ~2–3× baseline (Anitua/Labtician; correction from generic 1.6×) | Leukocyte-poor / leukocyte-free (P-PRP) | PDGF, TGF-β1, VEGF, EGF, HGF, IGF-1, FGF (+ fibrinogen/fibronectin) | CaCl₂ (PRGF Activator); thermal gelation (~76°C) for GEL | KMU10 injectable PRGF (~8 mL draw); KMU10-GEL ~3 mL autologous gel; KMU10-SCP super-concentrated | Yes (activated clot; GEL = autologous fibrin gel filler) | Facial rejuvenation, mesotherapy; GEL = autologous wrinkle filler (glabella, crow's feet, NLF, marionette), no HA/lidocaine | Yes — injectable PRGF usable scalp/AGA (GEL not for hair) | Yes (CE, tissue regeneration; TÜV-SÜD) | Not confirmed (no aesthetic-specific US clearance) | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (finales); ref KMU10; KMU10-SCP; KMU10-GEL | P-PRP (pure/LP) | Aesthetic: case series/pilot; PRGF platform: RCT-level | Moderate (aesthetic) / High (platform) | Autologous → donor variability; GEL is resorbable (not an HA filler); dental/ophthal/trauma ENDORET kits are non-aesthetic (see sheet 09) | Verified | https://medicalmix.com/productos/bti/kit-monouso-individual-obtencion-endoret-gel/?p=4974 | Distributor datasheet + BTI + AEMPS
ENDORET Dermatología / TOPIC LINE (KMU10-TPC) | BTI Biotechnology Institute SL | Vitoria-Gasteiz, Spain | PRGF-derived autologous TOPICAL serum (not injected) | Sodium citrate 3.8% (0.4 mL/9 mL tube) | BTI PRGF System V + Plasmaterm H (76°C gelation) | 580g × 8 min (RT) | 1 | ~2–3× baseline | Leukocyte-poor / leukocyte-free (P-PRP) | PDGF, TGF-β1, VEGF, EGF, HGF, IGF-1, FGF | CaCl₂-activated plasma gelated (76°C) then mixed 2:1 → cream/serum | 'Endoret-Serum' in airless pump; ~6-mo refrigerated; multi-week home course | No (topical serum) | Topical facial rejuvenation, sensitive skin, post-laser/energy-device aftercare, skin-barrier | Possibly topical scalp (not established) | Yes (CE, tissue regeneration) | Not confirmed | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (finales); ref KMU10-TPC | P-PRP-derived topical (LP) | Case series (Anitua topical serum) | Moderate | Topical only — do not inject; refrigeration/expiry; evidence limited to case series | Verified | https://pmc.ncbi.nlm.nih.gov/articles/PMC9528915/ | Peer-review (case series) + BTI + AEMPS
Tropocells PRP Kit | Tropocells (Bioregen S.r.l.) | Italy | PRP | ACD-A | Compatible standard centrifuge | ~900–1500g × 5–8 min (protocol not publicly standardized) | 1–2 | 3–5× estimated | LR or LP (kit-variant dependent — not clearly documented) | Standard PRP GF profile (estimated) | CaCl₂ activator included | 3–6 mL (estimated) | With activation | Facial rejuvenation, hair, mesotherapy (Italian aesthetics market) | Yes | CE marked (Italy/EU) | Not marketed in US | Available Spain/EU via distributors ¦ AEMPS CCPS 17-feb-2025: Tropocells PRP ref 11MACD7SGB — exposición mínima (finales) | Estimated L-PRP | Class III | Low (very limited published data; regional Italian brand) | Limited peer-reviewed evidence; spin protocol non-standardized; manufacturer data not independently verified; regional market brand | UNVERIFIED | https://www.tropocells.com/ | Manufacturer (unverified)
RegenKit-BCT-3 | Regen Lab SA | Le Mont-sur-Lausanne, Switzerland | PRP | Citrate + polymer gel separator | Standard bench centrifuge | 900g × 5 min | 1 | 1.5–2× (3 tubes → higher total platelet dose; per-mL concentration same as BCT) | LR (same as BCT) | Same as BCT — PDGF, TGF-β, VEGF, EGF | Same as BCT — CaCl₂ or THT thrombin | 8–12 mL combined (3× 3–4 mL tubes; from 3× 8 mL blood) | With activation | Higher volume facial rejuvenation, body (stretch marks, décolleté), scalp (larger coverage) | Yes — preferred for scalp coverage (higher volume) | Yes — CE | Registered | Spain: available ¦ AEMPS CCPS 17-feb-2025: RegenKit-BCT ref RK-BCT-T y RK-BCT-3 — exposición mínima | L-PRP | Class III | Moderate (same as BCT; higher volume = more complete treatment) | Requires 3 blood draws per treatment; coordinate collection timing | Verified | https://www.regenlab.com/products/ | Manufacturer
Cellenis Derma | Cellenis Medical Ltd | Israel / UK / Global | PRP | Sodium citrate + polymer gel | Standard centrifuge | 900g × 5 min | 1 | 2–3× baseline | LR | PDGF, TGF-β1, VEGF, EGF | CaCl₂ or thrombin | 3–4 mL (from 8–10 mL blood) | With activation | Facial rejuvenation, skin texture, periorbital, combination with micro-needling/laser | Limited dedicated scalp data | Yes — CE marked | FDA status unconfirmed | Available in Spain | L-PRP | Class III | Low-moderate | Modest platelet concentration (single spin); FDA status needs verification; limited published RCTs specific to Cellenis brand | Draft | https://www.cellenis.com/ | Manufacturer
Cellenis Hair | Cellenis Medical Ltd | Israel / UK / Global | PRP | Sodium citrate + polymer gel (hair-specific tube formulation) | Standard centrifuge | 900g × 5 min (Spin 1); 1500g × 3 min (Spin 2 — optional add-on) | 1–2 | 3–5× baseline (with optional Spin 2) | LR | PDGF, TGF-β1, VEGF (higher effective dose at scalp) | CaCl₂ or thrombin | 2–4 mL (hair-dedicated concentration) | With activation | Scalp: AGA (androgenetic alopecia), alopecia areata (off-label) | Primary indication | Yes — CE marked | FDA status unconfirmed | Available in Spain | L-PRP | Class III | Low-moderate | Spin 2 protocol less standardized; hair-specific differentiation from Derma line limited; FDA unconfirmed; insufficient head-to-head data vs EmCyte/Arthrex | Draft | https://www.cellenis.com/ | Manufacturer
Nahita-Jet PRP Kit | Nahita (Suministros Hospitalarios SA) | Spain | PRP | ACD-A (in pre-loaded kit tubes) | Compatible with standard centrifuge (Nahita offers branded centrifuges) | ~1500g × 5–8 min (1-spin recommended protocol) | 1 | 3–5× estimated | LR (single-spin; buffy coat included) | Standard PRP GF profile (estimated) | CaCl₂ or platelet activator included in kit | 4–6 mL (from 20 mL blood — estimated) | With activator | Hair (AGA), facial rejuvenation, aesthetic clinics (Spain-primary market) | Yes — widely used in Spanish aesthetic clinics for hair protocols | CE marked (Class IIb blood processing device) | Not primarily marketed for US | Spanish brand; AEMPS-registered/notified as medical device | L-PRP (estimated) | Class III (limited published data; mainly Spanish practitioner experience) | Low-moderate — regional brand with strong Spanish presence but limited peer-reviewed publications | Limited peer-reviewed data; regional brand; spin protocol less standardized than global brands; published evidence minimal; clinical data largely anecdotal | UNVERIFIED | https://www.nahita.com/ | Manufacturer (unverified — website not confirmed active)
Harvest SmartPrep 2 | Harvest Technologies (Terumo BCT) | Plymouth, MA, USA / Japan | PRP | ACD-A | Harvest SmartPrep 2 automated system | Automated 2-step centrifuge | 2 | 3–5× baseline | Variable (LP or LR adjustable) | Standard PRP GF profile | CaCl₂ or thrombin | Variable — 3–20 mL depending on protocol | With activation | Primarily orthopedic; facial rejuvenation (off-label aesthetic use) | Limited data | Yes — CE | FDA 510(k) cleared (K053001 and updates) | Available in Spain | L-PRP or P-PRP depending on protocol | Class II (orthopedic); Class III (aesthetic) | Low for aesthetic (primarily orthopedic literature; few aesthetic-specific publications) | Large expensive automated device — primarily hospital/surgery setting; aesthetics = off-label; not optimized for aesthetic volumes | Draft | https://www.harvesthpg.com/ | Manufacturer
RegenKit-BCT | Regen Lab SA | Le Mont-sur-Lausanne, Switzerland | PRP | Citrate anticoagulant + thixotropic polymer gel separator in tube | Standard bench centrifuge (any CE-marked) | 900g × 5 min | 1 | 1.5–2× baseline (limited by single-spin; polymer gel forms physical barrier) | LR — buffy coat retained above separator gel; leukocytes included (moderate WBC) | PDGF-AB, TGF-β1, VEGF, EGF, IGF-1 (standard PRP profile) | Optional: CaCl₂ or autologous thrombin (RegenKit-THT); may inject without activation | 3–4 mL (from 8 mL whole blood) | With activation → soft gel | Facial rejuvenation, skin quality, hair loss (AGA), mesotherapy supplement, post-filler/toxin regeneration | Yes — one of the most widely used in ES/EU hair protocols | Yes — CE certified (IVD-like autologous) | Registered as blood processing device (not drug) | Widely available Spain/EU; one of the most distributed PRP kits in Europe ¦ AEMPS CCPS 17-feb-2025: RegenKit-BCT ref RK-BCT-1-2 — técnica cerrada y exposición mínima | L-PRP: Leukocyte-rich, Low fibrin density | Class III (aesthetic); some Class II comparative studies for hair | Moderate (strong distribution = publication bias; solid clinical track record in EU) | Tube must not be overfilled; gel separator position critical; single spin gives lower platelet recovery than 2-spin; polymer gel cannot be reprocessed | Verified | https://www.regenlab.com/products/ | Manufacturer
Biomet Biologics GPS III | Zimmer Biomet | Warsaw, IN, USA | PRP | Sodium citrate | GPS III closed-loop centrifuge | 3200g × 15 min | 1 | 3–8× baseline | LR (single high-spin retains leukocytes and red cell fragments) | Standard LR-PRP profile — TGF-β1, PDGF, VEGF; also includes inflammatory cytokines (IL-1β from WBCs) | CaCl₂ or thrombin | ~3–5 mL | With activation | Primarily orthopedic; minimal aesthetic literature | Very limited | Yes — CE | FDA 510(k) cleared (K050932) | Available via distributors | L-PRP: LR, Low fibrin density | Class II (orthopedic only); Class IV (no aesthetic RCTs) | Very low for aesthetic (orthopedic-only evidence base) | Not designed for aesthetic use; LR profile with inflammatory cytokines may be suboptimal for periocular/sensitive zones; proprietary closed-loop system | Draft | https://www.zimmerbiomet.com/biologics | Manufacturer
ProGen PRP | ProGen Biologics | USA | PRP | Not confirmed | Not confirmed | Not confirmed | Not confirmed | Not confirmed | Not confirmed | Not confirmed | Not confirmed | Not confirmed | Not confirmed | Claimed: aesthetic PRP | Not confirmed | CE status unconfirmed | FDA status unconfirmed | Not available in Spain (US brand) | Not classified (insufficient data) | Class IV (insufficient data) | Not assessable | Brand name identified in user seed list; manufacturer details and product specifications could not be independently verified. List in gaps. | UNVERIFIED | No verified URL found | Not verified
EmCyte PurePRP II | EmCyte Corporation | Fort Myers, FL, USA | PRP | ACD-A | EmCyte PurePRP centrifuge | Spin 1: 400g × 10 min (remove RBCs); Spin 2: 1300g × 10 min (platelet concentrate) | 2 | 4–8× baseline (LP mode); 3–5× (LR mode — draw technique dependent) | Selectable LP or LR depending on draw — unique feature | PDGF-AA/AB/BB, TGF-β1/β2, VEGF, EGF, bFGF, IGF-1 (full growth factor profile at high concentration) | Optional CaCl₂ or calcium gluconate; many protocols uninactivated | 2–6 mL LP / 3–8 mL LR (from 30–60 mL blood) | With activation | Hair (AGA — primary commercial focus), facial rejuvenation, acne scars, under-eye, body (stretch marks), combined with micro-needling | Yes — one of highest-evidence commercial systems for AGA (multiple published RCTs) | Yes — CE marked | FDA 510(k) cleared (K192673) | Available in Spain | P-PRP (LP mode) or L-PRP (LR mode) — clinician-selectable | Class II for hair (multiple RCTs); Class III for facial | High for hair; Moderate for facial (best-published commercial aesthetic PRP system for AGA) | Large blood draw (30–60 mL) may deter some patients; 2-spin time ~20 min additional chair time; proprietary centrifuge required; higher upfront kit cost; selection of LP vs LR important — not interchangeable | Verified | https://www.emcyte.com/pure-prp/ | Manufacturer
Eclipse PRP | Eclipse Aesthetics LLC | Irving, TX, USA | PRP | ACD-A | Eclipse centrifuge | 1500g × 5 min × 2 | 2 | 4–6× baseline | LP (2-spin system preferentially depletes leukocytes) | PDGF, TGF-β1, VEGF, EGF | CaCl₂ or autologous thrombin | 4–6 mL (from 22 mL blood) | With activation | Vampire Facelift® (Runels registered trademark — Eclipse specifically named), hair (Vampire Hair Restoration®), facial rejuvenation, acne scars, skin texture | Yes — Eclipse protocol integral to published Vampire Hair studies | Yes — CE | FDA registered | Available in Spain/EU via distributors | P-PRP: LP, Low fibrin density | Class III (mostly Eclipse-sponsored or Runels-branded studies) | Moderate (high market penetration; Vampire Facelift brand recognition; LP reduces pain/bruising claim; publication quality variable) | 'Vampire' branding = marketing; most evidence is weak/sponsored; proprietary centrifuge required; 2-spin protocol more complex; cost per treatment higher | Draft | https://www.eclipseaesthetics.com/eclipse-prp/ | Manufacturer
EmCyte GenesisCS (GS-PRP) | EmCyte Corporation | Fort Myers, FL, USA | PRP | ACD-A | EmCyte GenesisCS centrifuge | 2-stage: 400g × 10 min + 1300g × 10 min | 2 | 6–10× baseline (super-concentrated option) | LP (ultra-filtered) | High-concentration GF profile (same profile as PurePRP but higher yield) | Optional CaCl₂ | 2–4 mL ultra-concentrated (from 60–120 mL blood) | With activation | Hair restoration (high-density protocol), large-area body, hyperdiluted PRP | Yes — ultra-concentrated scalp protocols | Yes — CE | FDA 510(k) cleared | Available in Spain | P-PRP: LP, ultra-concentrated | Class III (limited studies at ultra-concentration) | Low-moderate (very large blood draw; concentration benefits over standard PurePRP not clearly proven) | Very large blood draw (up to 120 mL) — not routine for all patients; marginal evidence that ultra-concentration outperforms standard; proprietary large-volume system | Draft | https://www.emcyte.com/genesis-cs/ | Manufacturer
Arthrex ACP (Autologous Conditioned Plasma) | Arthrex Inc | Naples, FL, USA | PRP | ACD-A (Acid Citrate Dextrose Solution A) | Proprietary double-syringe centrifuge system (ACP centrifuge) | 1500g × 5 min | 1 | 2–4× baseline | LP (double-syringe design physically separates PRP from buffy coat — key design feature for LP PRP) | PDGF-AB, TGF-β1, VEGF, IGF-1, HGF — LP profile avoids pro-inflammatory IL-1β from leukocytes | Optional CaCl₂; many protocols inject without activation (ACP philosophy) | 3–4 mL (from 15 mL blood) | With activation (optional) | Hair (AGA), facial rejuvenation, periorbital, acne scars, post-procedure, orthopedic (primary use) | Yes — used in AGA hair loss protocols with supporting literature | Yes — CE marked | FDA 510(k) cleared (K070834 — blood processing for orthopedic use) | Available in Spain ¦ AEMPS CCPS 17-feb-2025: 'Jeringa Doble ACP' ref ABS-10014 — técnica cerrada y exposición mínima | P-PRP: Pure (LP), Low fibrin density | Class II (orthopedic); Class III (aesthetic) | Moderate-high for hair loss; moderate for facial (LP PRP: less inflammatory response, possibly better tolerated peri-orbitally) | FDA clearance is orthopedic; aesthetic use is off-label; lower platelet yield than 2-spin systems; proprietary centrifuge required | Draft | https://www.arthrex.com/biologics/acp-double-syringe-system | Manufacturer
RichPlat Ortho | Fidia Pharma USA Inc. (parent: Fidia Farmaceutici, Italy) | USA / IT | PRP (autologous, closed system) | ACD-A 2 mL | BOOST 2 Flex (pre-programmed) | 600g (2000 rpm) × 8 min | 1 | 2.3× over baseline | Leukocyte profile not specified; 99% RBC reduction (likely LP) | PDGF/TGF-β/VEGF (generic PRP — not product-specific) | CaCl2 10% 1 mL (optional) | 8 | No (liquid PRP) | Marketed orthopedic (bone-graft mixing); aesthetic use would be off-label | Not indicated (orthopedic registration) | TBD (US product) | Rx Only | TBD (Fidia is EU — verify AEMPS) | Low–moderate (2.3× conc, leukocyte-poor, no default activation) | Limited (orthopedic) | Draft | Orthopedic indication; aesthetic/scalp use off-label; an aesthetic 'RichPlat' line (Fidia Italy) likely exists — verify separately | Verified (manufacturer) | https://www.fidiaprp.com/ | Manufacturer
RegenACR Classic (A-PRP) | Regen Lab SA | Switzerland (Le Mont-sur-Lausanne) | PRP (leukocyte-poor / Pure) liquid | Sodium citrate (pharma grade; reversible) | Swing-out/fixed-angle; RegenPRP Centri (≤2200 RCF) | 1500 g × 5 min | 1 | 1.6× native (recovery >80%) | Leukocyte-poor (~96.7% granulocyte depletion; RBC >99.7% depleted) | PDGF, TGF-β, VEGF, EGF, FGF, IGF (+ plasma IGF/HGF, fibrinogen) | None dedicated (in-situ tissue activation) | ~5–6 mL PRP (1 BCT tube; 10 mL draw) | No | Facial/skin rejuvenation (single/small area), collagen induction, acne & keloid scars (adjunct) | Yes — scalp/AGA (general PRP evidence) | Yes, Class IIb, CE 2797; ISO 13485 + MDSAP | Regen PRP components 510(k)-cleared for PRP prep; aesthetic use off-label | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — técnica cerrada; ref R-ACR C1/B | P-PRP (LP; DEPA dose/purity derived — not published) | Moderate for facial skin (Abuaf 2016; Redaelli 2010); Moderate–High AGA | Media–alta | Use within 4 h; single-use; do not exceed 2200 RCF; LP profile; single tube = lowest platelet dose | Verified | https://estheticmed.ch/wp-content/uploads/2022/01/Doc-Regen-ACR-A-PRP-ENG-pp.pdf | Manufacturer brochure
RegenACR Extra (A-PRP) | Regen Lab SA | Switzerland | PRP (leukocyte-poor / Pure) liquid | Sodium citrate (pharma grade; reversible) | Swing-out/fixed-angle; RegenPRP Centri (≤2200 RCF) | 1500 g × 5 min | 1 | 1.6× native (recovery >80%) | Leukocyte-poor (~96.7% granulocyte depletion) | PDGF, TGF-β, VEGF, EGF, FGF, IGF | None dedicated (in-situ; some regional bundles add 10% calcium gluconate — optional) | ~10–12 mL PRP (2 BCT tubes; ~20 mL draw) | No | Facial + neck rejuvenation, larger/higher-volume areas, skin texture/elasticity, scars | Yes — scalp/AGA (larger volume favours multi-area/hair) | Yes, Class IIb, CE 2797; ISO 13485 + MDSAP | Regen PRP 510(k) for prep; aesthetic off-label | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — técnica cerrada (fase 4a/5) / exposición mínima (fase 4b); ref R-ACR C2/B | P-PRP (LP; DEPA derived) | Moderate facial/skin; Moderate–High AGA | Media–alta | Use within 4 h; ≤2200 RCF; 2-tube pooling = the minimal-exposure step flagged by AEMPS | Verified | https://estheticmed.ch/wp-content/uploads/2022/01/Doc-Regen-ACR-A-PRP-ENG-pp.pdf | Manufacturer brochure
DIVES MED PRP COMPLETE SOLUTION | T-Biyoteknoloji (OEM; brand Dives Med) | Turkey (OEM) / marketed DE-PL | PRP (likely P-PRP/LP) | Sodium citrate 1 mL/10 mL tube (2 tubes/kit) | Standard centrifuge (not specified) | Not confirmed (T-Biyoteknoloji house family — likely 500 RCF×3/1150×4 or ~830g) | 1 (typical; not confirmed) | Not confirmed | Not confirmed (likely LP) | PDGF, TGF-β, VEGF, EGF, IGF-1 (generic) | None default (autologous) | Not confirmed (~4 mL/procedure; 2 procedures/kit) | No | Wrinkles, scars, stretch marks, facial rejuvenation | Yes — hair growth | Yes, CE Class IIb (2195) | Not confirmed | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima; ref 4251871102881 | Not classifiable (no published dose/purity) | Class IV (regional brand; no independent RCT) | Low | Brand-specific spin protocol not published; OEM = T-Biyoteknoloji inferred (CE 2195 + Na-citrate match); kit incl. 20G/30G/32G needles + 5 mL syringes | Draft | https://fillerpharma.com/dives-med-prp-complete-solution-1-set/lang:EN | Distributor spec sheet (no mfr IFU located)
RegenACR Plus (A-PRP + autologous fibrin gel) | Regen Lab SA | Switzerland | PRP (LP) + autologous thrombin serum → autologous fibrin gel | Sodium citrate (BCT tubes) + autologous thrombin (ATS tube) + calcium gluconate at mixing | Swing-out/fixed-angle; RegenPRP Centri (≤2200 RCF) | 1500 g × 5 min (BCT and ATS tubes) | 1 | 1.6× native for PRP fraction (recovery >80%) | Leukocyte-poor (~96.7% granulocyte depletion) | PDGF, TGF-β, VEGF, EGF, FGF, IGF | Exogenous — autologous thrombin serum (RegenATS) + calcium → slow-release fibrin gel | ~10–12 mL PRP (2 BCT) + fibrin-gel fraction (1 ATS); ~20–24 mL draw | Yes — autologous fibrin gel (sustained GF release; volumising) | Volume restoration, deep wrinkle fill, biostimulation, combination facial rejuvenation | Yes (gel less typical for diffuse scalp) | Yes, Class IIb, CE 2797; ISO 13485 + MDSAP | Regen PRP 510(k) for prep; aesthetic off-label | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (etapas finales); ref R-ACR C/BA | P-PRP + exogenous activation (DEPA derived) | Moderate (facial PRP) + mechanistic support for fibrin gel; PRP+fibrin evidence lower | Media | Use within 4 h; final PRP+ATS mixing = the minimal-exposure step; apply gel promptly | Verified | https://estheticmed.ch/wp-content/uploads/2022/01/Doc-Regen-ACR-A-PRP-ENG-pp.pdf | Manufacturer brochure
CUREVITAL PRP (CU500) | T-Biyoteknoloji (ES distr. IBOR Médica) | Turkey (Bursa) | PRP (P-PRP / pure) | Not confirmed in IFU (sodium citrate presumed) | Standard centrifuge | 1500 g × 4 min (single spin per CU500 IFU) | 1 | Not confirmed | Not confirmed (single low-g → likely LP) | PDGF, TGF-β, VEGF, EGF, IGF-1 (generic) | None in IFU (autologous) | ~3–4 mL PRP from 10 mL blood | No | Facial/skin rejuvenation (generic PRP) | Yes (generic PRP hair use) | Yes, CE Class IIb (Szutest 2195); EO-sterile, 5-yr shelf | No 510(k) found | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (iniciales y finales); ref CU500 | Not classifiable (no published dose/purity) | Class IV (regional brand; no independent RCT) | Low | Tube-only (AEMPS: no incluye agujas); anticoagulant type not stated in IFU; single 1500g spin = simpler/lower-control | Verified | http://ifuport.com/cu500/ | Manufacturer IFU + ES distributor
Dr.PRP (DrPRP; 20 cc / 30 cc) | Rmedica Co. Ltd. (US distributor: Dr. PRP USA LLC, Stafford TX) | Rep. de Corea (Seoul) — US distribution | PRP — double-compartment dual-centrifugation closed system (double seal) | ACD-A (kit completo — includes anticoagulant) | Standard bench centrifuge | Not confirmed (g×min not published; dual-spin) | 2 (dual centrifugation — corrects prior single-spin gel-separator guess) | Not confirmed ('high-concentration' claimed; prior 1.5–2× likely understated for a double-spin) | LR (buffy coat; adjustable) | PDGF, TGF-β, VEGF, EGF, IGF (generic) | Optional CaCl₂/thrombin | Not confirmed (20 cc / 30 cc versions) | No | Facial rejuvenation, skin texture | Yes — hair (AGA) | Yes — CE (MDD/MDR) | FDA 510(k) CLEARED (CBER): BK230818 (20 cc; 2023) + BK251201 (30 cc) | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima; ref DrPRP | L-PRP | Class IV (regional brand; no independent RCT — but FDA-cleared) | Moderate (upgraded — dual-spin closed system with FDA 510(k) + CE) | CORRECCIÓN: fabricante es Rmedica (no AHC); dispositivo de doble centrifugación (no tubo con gel de una sola vuelta); FDA ahora confirmado como cleared | Verified | https://www.fda.gov/vaccines-blood-biologics/substantially-equivalent-510k-device-information/bk230818-dr-prp | Regulator (FDA) + Manufacturer
Next PRP (N500 / 94223N500) | T-Biyoteknoloji (brand Next PRP / T-LAB Europe) | Turkey (Bursa) | PRP — Pure P-PRP (LP; double-spin option) | Optional — anticoagulant-free OR sodium citrate | Standard centrifuge | 1500 g × 4 min (P01/P02); 8 min (P03); optional double-spin | 1 (single) or 2 (double option) | 2–3× dose (mfr); 98% purity; 90% recovery (mfr) | LP-PRP (Pure >90% platelets) | PDGF, TGF-β, VEGF, EGF, IGF-1 (generic) | None default (autologous; anticoagulant-free protocol available) | ~3–4 mL from 10 mL draw | No | Facial rejuvenation, skin quality | Yes — scalp | Yes, CE Class IIb (Szutest 2195); closed system | No 510(k) found | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima; ref N500 | Not classifiable (dose not independently published; high purity per mfr) | Class IV (regional brand; no independent RCT) | Low | Purity/recovery/2–3× are manufacturer claims; 10 mL syringe-to-tube format; verify anticoagulant choice per protocol | Verified | http://ifuport.com/n500/ | Manufacturer IFU
PRPHD PRP KIT (HD1000) | T-Biyoteknoloji (T-LAB private label; ES distr. IBOR Médica) | Turkey (Bursa) | PRP — selectable LP (Pure) / LR (Amber) | Buffered sodium citrate 3.2% (1 mL/10 mL tube); gel-free | Any lab centrifuge; swing-out recommended | Pure 500 RCF×3/1150×4; Amber 500×6/1150×8; TR sheet cites ~830g (~2300–2500 rpm) ×8 swing | 1 per tube | Mfr claim up to 3.5 M plaquetas/µL (≈ up to ~10–14×; marketing) | Selectable LP (Pure) / LR (Amber) | PDGF, TGF-β, VEGF, EGF, IGF-1 (generic) | None default (autologous) | ~4 mL (2 mL/tube; 20 mL draw) | No | Facial rejuvenation, skin quality, scars/striae | Yes — scalp/AGA | Yes, CE Class IIb (Szutest 2195) | No 510(k) found | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima; ref HD1000 | Not classifiable (dose not independently published; high purity) | Class IV (regional brand; no independent RCT) | Low | 3.5 M/µL is a manufacturer claim not independently verified; gel-free needs careful buffy aspiration | Verified | http://ifuport.com/hd1000/ | Manufacturer IFU + ES distributor
KIT PRP T-LAB (T-1000) | T-Biyoteknoloji Laboratuvar (brand T-LAB) | Turkey (Bursa) | PRP — selectable P-PRP (LP) / L-PRP (LR) | Buffered sodium citrate 3.2% (1 mL/10 mL tube); gel-free | Any lab centrifuge (T-LAB unit available) | Pure 500 RCF×3 min (fixed) / 1150 RCF×4 min (swing); Amber 500×6 / 1150×8 (mfr also cites ~830g) | 1 per tube (optional 2-step for Amber) | Not stated as ×; mfr 80–95% recovery / 95–99% purity | Selectable LP (Pure) / LR (Amber buffy-coat) | PDGF, TGF-β, VEGF, EGF, IGF-1 (generic) | None default (optional CaCl₂) | ~4 mL (from 20 mL blood) | No | Facial rejuvenation, skin quality, periorbital, microneedling adjunct, scars/striae | Yes — AGA | Yes, CE Class IIb (Szutest 2195) | No US 510(k) found | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima; ref T-1000 | Not classifiable (dose not published; high-purity P-PRP) | Class IV (regional brand; no independent RCT) — most-established of the T-LAB group | Low–Moderate | Operator-dependent yield; gel-free needs careful buffy aspiration; mfr % claims not independently verified | Verified | http://ifuport.com/t1000/ | Manufacturer IFU
S&M PRP MONOCAP | Str Biyo Teknolojileri Medikal Ltd. | Turkey (Çorum) | Single-cap ('MONOCAP') closed-system PRP tube (draw + PRP extraction from same rubber cap); gel or gel-free | Included (kit completo); type not confirmed (citrate typical) | LABFUGE / standard centrifuge | Not confirmed | Not confirmed (single typical) | Not confirmed | Not confirmed | PDGF, TGF-β, VEGF, EGF, IGF-1 | None routine | Not confirmed (10–12 mL draw) | No | Aesthetics, regenerative medicine | Yes (hair restoration) | Yes (CE MDR 2017/745; ISO Class 5 cleanroom/GMP) | Mfr claims 'FDA approved' — UNVERIFIED (likely establishment registration not clearance) | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (finales); ref STR-PRPKIT-01 | Not classified | Class IV (no independent RCT) | Low | 'FDA approved' marketing claim unverified; single-cap = draw + extraction share one port | Draft | https://www.strmedical.com/en/sm-prp-kit-monocap | Manufacturer
REIN PRP | Kaizen Endüstri ve Sanayi Dis Ticaret Ltd. | Turkey (Balıkesir) | Single-spin PRP kit/tube; gel-free | Sodium citrate 3.13% 1 mL/10 mL tube (INCLUDED — kit completo) | Standard centrifuge | Not confirmed (single typical) | 1 (typical) | Not confirmed (mfr claims 98–99% 'purity') | Not confirmed | PDGF, TGF-β, VEGF, EGF, IGF-1 | None routine | Not confirmed (~4–5 mL from 10 mL tube) | No | Dermatology/aesthetic, wound healing | Yes (hair) | Yes (CE Class IIb) | Not confirmed | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (finales); ref OM-RE1-0001 | Not classified | Class IV (no independent RCT) | Low | Specs are manufacturer marketing only; spin/concentration unpublished | Draft | https://www.kaizenendustri.com/en/rein-prp-nedir | Manufacturer
Xerthra PRP KIT | Biovico Sp. z o.o. | Poland (Gdynia) | Single-spin PRP separation device (tube-based); PURE variant = leukocyte-reduced | Sodium citrate 3.13% (NOT in kit — user adds; study used 1.5 mL/13.5 mL blood) | Dedicated centrifuge | 3500 rpm × 5 min | 1 | Up to 9.7× (final >1,000,000 PLT/µL) | Standard = LR; Xerthra PURE = leukocyte-reduced (LP) | PDGF, TGF-β, VEGF, EGF, IGF-1 | None routine | Not confirmed (~4–6 mL) | No | Facial skin rejuvenation, biostimulation | Yes (androgenetic hair) | Yes (CE-marked) | Not confirmed | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (iniciales y finales); ref PRO00447 | Not classified | Class IV (one preliminary comparative study) | Low | Anticoagulant not supplied — operator must add correct sodium citrate | Verified | https://biovico.com/en/products/xerthra-prp-kit-2 | Manufacturer + peer-review comparison
Accelerate (EmCyte) | EmCyte Corporation | USA (Fort Myers, FL) | Single-spin whole-blood PRP kit (simpler entry line vs double-spin PurePRP SP) | ACD-A | EmCyte centrifuge (Executive/Sapphire/PurePRP) | Single spin; exact g×min not confirmed for 'Accelerate' | 1 (vs 2 for PurePRP SP) | ~2–5× (single-spin ceiling; EmCyte single-spin measured ~4.8× — far below PurePRP double-spin 7–15×) | Not confirmed for Accelerate (single-spin buffy-coat typically LR) | PDGF, TGF-β, VEGF, EGF, IGF-1, bFGF | None routine (optional CaCl₂) | ~7–8 mL from ~20 mL draw (line-dependent; not confirmed for Accelerate) | No | Facial rejuvenation, skin quality | Yes (androgenetic hair) | Yes (implied — CE-marked EU device) | US mfr; marketed under FDA device registration (not a PRP-specific 510(k)) | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — con exposición ambiental mínima (iniciales); ref 620-12-02; 620-22-06; 620-22-08 | Not classified (DEPA not reported for Accelerate) | Class III–IV (PurePRP family has literature; 'Accelerate' itself sparse) | Moderate (family) / Low (this kit) | Single-spin gives lower dose than PurePRP double-spin — do not assume PurePRP concentrations | Draft | https://emcyte.com/ | Manufacturer + peer-review device review
SBL.PRP.3 15 ml | D-Biotech SRL (mfr, Italy); brand Solutions Biomedical Lab (Barcelona) | Italy / Spain | Tube+gel — 15 mL borosilicate U-bottom tube; thixotropic separator gel; closed circuit; lateral-port transfer | ACD-A (pre-loaded in tube) | Any lab centrifuge (SBL supplies compatible units) | ~6 min (exact g/rpm not published) | 1 | 5.4× baseline; ~80% recovery; RBC contamination ~0.28 (low) | Leukocyte-poor ('pure'); +M variant adds monocytes | PDGF, TGF-β, VEGF, EGF, IGF-1 (generic) | None default (optional CaCl₂) | ~4 mL concentrated PRP (of ~7.5 mL plasma from 15 mL tube) | No | Aesthetics, dermatology, trichology, gynecology, orthopedics, urology, dental | Yes (trichology / hair restoration) | Yes — CE 0425 Class IIa | Not confirmed | En listado AEMPS de sistemas PRP CERRADOS (CCPS, act. 17-feb-2025) — técnica cerrada / exposición mínima (finales); ref 'SBLxyzz' (placeholder; product = SBL.PRP.3) | High purity ~5.4× no activation (DEPA not formally published) | Class III–IV (no independent RCT located) | Low | rpm/g not stated; do not invert tube 2 min post-spin; blood-draw system not included | Verified | https://pdf.medicalexpo.es/pdf/solutions-biomedical-lab/sblprp3/4588938-283482.html | Manufacturer datasheet + AEMPS
Mybloc (PRP brand) | POSSIBLE BRAND CONFUSION — see Notes | Unknown | UNKNOWN — possible confusion | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Unknown | Unknown | Unknown | Unknown | N/A — cannot assess | Unknown | Not assessable | CRITICAL FLAG: 'Myobloc' (also 'Mybloc') is the brand name for rimabotulinumtoxinB (Ipsen/Solstice Neurosciences), a botulinum toxin type B product (neurotoxin). No verified PRP system named 'Mybloc' was found. This entry is flagged for user clarification — if a PRP kit by this name exists, it is undocumented in peer-reviewed literature. | UNVERIFIED | N/A — no PRP product found | Not verified — possible brand confusion with botulinum toxin

## Indication x Product
DRAFT — Indication × Product matrix for PRP/Autologous regenerative systems. Evidence grades: A = ≥2 RCTs; B = 1 RCT or multiple cohorts; C = case series/expert opinion; D = insufficient data / theoretical; ✗ = not indicated / contraindicated. Leukocyte preference (LP vs LR) noted where evidence exists.
Indication / Zone | LP-PRP (Arthrex ACP / Eclipse / Selphyl PRFM) | LR-PRP (EmCyte PurePRP / RegenKit-BCT / Dr.PRP) | PRGF-Endoret (BTI) | i-PRF (Miron/Choukroun) | L-PRF / A-PRF (membrane) | CGF — Medifuge (Silfradent) | Plasmagel (thermal filler) | Combined PRP + HA filler | Combined PRP + Biostimulator | Notes / Preferred Protocol
Acne Scars (atrophic — icepick, boxcar, rolling) | B — LP-PRP for acne scars: multiple RCTs (especially LP-PRP + micro-needling combination); Eclipse + Dermapen studies numerous | B — LR-PRP also well-studied for acne scars; PDGF from WBC may augment fibroblast activation; EmCyte studies published; LR not clearly superior to LP | C — PRGF acne scar (BTI limited publications for aesthetic acne; reasonable based on GF profile) | C — i-PRF + micro-needling for acne scars: emerging data; small number of studies | C — PRF membrane packed into punch excision sites or PRF serum post-dermabrasion | C — CGF serum + micro-needling; Italian case reports | ✗ — Not appropriate for acne scars (no regenerative effect + heat-denatured proteins not suitable for scar remodelling) | B — PRP + SubQ/hyperdiluted HA (Radiesse or soft HA) injection for rolling scars: volumetric + regenerative combination | C — PRP + PLLA (Sculptra) for widespread atrophic scarring: combined volume + collagen stimulus | Best evidence: LP-PRP + micro-needling (Dermapen/RF-microneedling) = Class A evidence (multiple RCTs). Protocol: 3 sessions × 4 weeks apart. PRP applied topically (micro-channelling) + injected subdermally at scar base. Subcision first for fibrous anchoring. Combine with 1–2 sessions HA subliminal for rolling scars.
Body — Stretch Marks (Striae distensae) | C — LP-PRP intradermal injection along striae: small series; promising | C — LR-PRP: RegenKit-BCT-3 large volume body protocol; modest evidence | C — PRGF body (limited published) | C — i-PRF body injection: emerging | ✗ | ✗ | ✗ — Thermal gel in striae: not indicated | C — PRP + micro-needling (body) for striae: some published series | C — PRP + PLLA (Sculptra body) or + HA skin booster for striae | Striae: PRP addresses colour (red→white) and partial texture improvement. Micro-needling combination essential. Multiple sessions (4–6) required. Evidence weaker than for facial/hair. Consider RF-microneedling + PRP for best outcomes.
Combined with Biostimulators (PLLA / CaHA / PCL) | C — LP-PRP + Sculptra (PLLA): inject PLLA first (deep); PRP (intradermal) after 14 days; not same session | C — LR-PRP + Radiesse (CaHA): emerging; some practitioners combine; separate sessions preferred | C — PRGF + biostimulator: reported in some protocols | C — i-PRF + biostimulator: theoretical | ✗ | ✗ | ✗ | C — HA + PRP ± biostimulator triple layered protocol: complex but published | C — This IS this cell: biostimulator + PRP = synergistic collagen induction | Rationale: PRP = early GF-mediated fibroblast stimulation (weeks); PLLA/CaHA = sustained scaffold stimulation (months). Temporal sequencing important: do PRP course first (3× monthly) THEN biostimulator, OR PRP same day as diluted biostimulator (deep CaHA + superficial PRP).
Combined with HA Filler (same session) | B — LP-PRP mixed into HA syringe (1:4 ratio PRP:HA) or injected in same zone separate layer; multiple published protocols | B — LR-PRP same session as HA: possible; deep HA bolus (periosteal) + superficial LR-PRP (intradermal) layered approach | C — PRGF + HA: some BTI centre publications | C — i-PRF + HA: theoretical (fibrin + HA scaffold); limited evidence | ✗ | C — CGF serum post-filler layered technique | C — Plasmagel AS the filler + HA: not standard; not recommended in same zone (unpredictable interaction) | A — This IS the combination; HA (structural base) + PRP (regenerative layer) = standard aesthetic protocol; evidence strongest for this dual approach | C — HA + PRP + Polynucleotide triple combination: emerging | Protocol: HA deep structural filler FIRST; wait 3–5 min; then PRP intradermal over same zone (separate layer). Or: premix 0.5 mL PRP + 0.5 mL HA 20 mg/mL → inject as combined bolus. Avoid premixing LP-PRP with cross-linked HA (activation may compromise HA rheology). Non-cross-linked HA skin booster + PRP = synergistic; cross-linked HA + PRP = inject in separate layers.
Face Rejuvenation — Global / Skin Quality (texture, tone, pore size, luminosity) | B — Multiple small RCTs for facial rejuvenation; LP preferred for sensitive skin (less erythema/swelling); Selphyl PRFM small case series | B — LR-PRP has more growth factor load (especially PDGF from monocytes/thrombocytes); may produce stronger collagen stimulus at cost of more post-procedure swelling; RegenKit-BCT widely used in EU practice | B — PRGF Endoret facial series (Anitua/BTI publications); lower GF yield but LP profile favoured by users who see better skin quality vs LR | C — i-PRF mesotherapy for skin quality; 0.05–0.1 mL per injection point; technique-sensitive; evidence emerging | C — PRF liquid serum used in mesotherapy; limited published data for facial rejuvenation (primary use = wound) | C — Italian case series for facial skin quality; CGF injectable serum; limited | ✗ — Not a growth factor approach; purely volumetric if used in face | A — PRP + HA filler combination is CLASS-A evidence in periorbital/NLF zone; PRP activates HA scaffold (synergistic collagen stimulation claimed) | B — PRP + Profhilo (published protocols) or PRP + PLLA (theoretically synergistic) | Preferred for face: LP-PRP or PRGF injected intradermally with 30G cannula / needle; 1–2 sessions/year; combine with HA filler or biostimulator for structural + regenerative dual effect. Evidence is for general skin quality — disease-specific claims (wrinkles) have weaker evidence.
Periorbital / Tear Trough | B — LP strongly preferred periorbital (less bruising, less inflammatory cytokines); Selphyl PRFM published for tear trough (case series + small RCT); small volume suits this zone | C — LR periorbital acceptable if LP unavailable; higher bruising/swelling risk; 0.1–0.2 mL per side maximum | C — PRGF LP profile well suited periorbital; some Anitua series; reasonable choice | C — i-PRF small volumes (0.1–0.2 mL) periorbital; some published tear trough cases; promising | ✗ — Membrane not applicable; PRF serum micro-injection possible but volume insufficient | C — CGF serum injection (very small vol) periorbital; case reports | ✗ — Periorbital zone: avoid Plasmagel (unpredictable gelation, risk of lumpiness in thin skin) | A — PRP + HA (tear trough filler, hypo-diluted) = well-established; PRP mixed into HA syringe or layered over HA | C — PRP + PN (polynucleotides) for periorbital hydration + regeneration; emerging | Tear trough: volume required < 0.5 mL per side. LP-PRP or PRGF preferred. Often combined with small volume HA (Restylane EyeLight, Teosyal Redensity II) for structural + regenerative dual effect. Avoid deep fat compartment injection with PRP — superficial subdermal/intradermal plane only.
Post-procedure Enhancement (post-laser / post-peel / post-RF / post-filler) | B — LP-PRP or PRGF topically + intradermally immediately post-ablative laser (fractional CO2 / erbium) = well-established combination; accelerates healing and enhances collagen remodelling | B — LR-PRP topical (micro-channelled / applied to ablated skin) also effective; slightly more downtime | B — PRGF widely used in BTI centres post-ablative procedures; LP profile especially valued for tolerability | C — i-PRF post-laser: topical application; case reports | C — L-PRF membrane applied to post-ablative skin (full face); primary data dental but extrapolated to aesthetic | ✗ — Injectable in post-laser setting | ✗ — Not post-procedure | C — PRP applied during filler session (no filler in same zone on same day for intradermal filler; same session OK for deep bolus + surface PRP) | C — PRP post-biostimulator (day 14–21 synergy) | Standard: Apply PRP/PRGF immediately post-fractional ablative laser to micro-channelled skin (topical + intradermal). Accelerates re-epithelialisation. Reduces erythema duration 30–50% (RCT evidence). PRP + Microneedling post-peel: wait until skin barrier restored (48–72 h). PRGF specifically well-documented for this (Anitua series).
Scalp & Hair — Androgenetic Alopecia (AGA) ¦ STRONGEST evidence in aesthetics | B — LP-PRP shows good results in AGA (Arthrex ACP studies; less inflammatory scalp response claimed); INJECT intradermally / subdermal scalp at 1 cm intervals; 4–6 sessions monthly → maintenance q3–6 months | A — BEST commercial evidence: EmCyte PurePRP LP-mode has multiple RCTs; LR-PRP also effective but potentially more scalp soreness; RegenKit BCT-3 used widely in EU; standard of care for PRP-based AGA | C — PRGF used in hair loss (Anitua et al. published series); LP; smaller evidence base vs commercial PRP systems | C — Emerging; i-PRF + PRP combination gaining traction (Miron RJ hair studies); small volume limits full scalp coverage; adjunct not monotherapy | C — PRF serum injected; membrane squeezed onto scalp post-hair transplant; A-PRF membrane on donor site; NOT standard AGA protocol | C — CGF serum used in Italian hair transplant centres; wound coverage post-FUE | ✗ — No GFs (heat destroys); inappropriate for hair loss | B — PRP + low-concentration HA (skin booster dose) published for AGA (Regenera/Hairomax combination); PRP + Juvederm/Restylane Vital — emerging | C — PRP + Polynucleotides (e.g. Plinest) synergy proposed; limited evidence; PRP + Profhilo for scalp hydration | STRONGEST evidence area for PRP. Preferred: LP-PRP or LP-mode EmCyte PurePRP; 4–6 sessions × 1 mL/cm² scalp coverage; monthly × 3 then quarterly maintenance. Combination with minoxidil / finasteride strongly augments results. Measure: Hamilton-Norwood scale + dermoscopy (FU density) at baseline and 6 months.
Under-Eye Dark Circles / Hollowing | C — LP-PRP small series for periorbital dark circles; anti-inflammatory LP profile reduces haemosiderin-related darkening claim; limited evidence | C — LR limited for true dark circles (more inflammation = risk of post-inflammatory darkening, especially Fitzpatrick III–VI) | C — PRGF periorbital published (Anitua/BTI); reasonable | C — i-PRF periorbital: 0.1 mL per side; promising technique | ✗ | ✗ | ✗ | B — HA tear trough filler (Restylane EyeLight / Teosyal Redensity II / Belotero) + PRP = standard approach for combined hollowing + dark circles | C — PN (Plinest/Nucleofill) periorbital + PRP combination | True dark circles = pigmentary + vascular + structural (volume) components. PRP helps structural/pigmentary component only; vascular component (periorbital vessels showing through thin skin) = address with stimulating collagen thickening (PRP + PN). HA filler for hollowing. Combine all three for complete result.

## Classification
DRAFT — PRP Classification Systems: DEPA (2009) / PAW (2012) / MARSPILL (2019). Purpose: standardise communication across studies and clinical reporting so 'PRP' is never used without qualifier. Sources: Dohan Ehrenfest 2009; Mishra 2012; Fadadu 2019; Kon E 2021 review.
Classification System | Year Proposed | Authors / Reference | Variables Classified | Variable Definitions | How to Apply — Example | Key Products Mapped | Limitation | Source Verification
CGF / Plasmagel |  |  | CGF: Class III (Italian case series; no RCTs). Plasmagel: Class III (case series); NOT regenerative — purely volumetric; GFs denatured by heat.
DEPA Classification | 2009 | Dohan Ehrenfest DM et al. — Classification of platelet concentrates: from pure platelet-rich plasma (P-PRP) to leucocyte- and platelet-rich fibrin (L-PRF). Trends Biotechnol. 2009;27(3):158–167 | 2 variables: (1) Leukocyte content (Pure = LP vs Leukocyte-rich = LR); (2) Fibrin architecture (low density = liquid PRP vs high density = fibrin matrix) | P-PRP (Pure Platelet-Rich Plasma): LP + Liquid [low fibrin]. L-PRP (Leukocyte-rich PRP): LR + Liquid [low fibrin]. P-PRF (Pure Platelet-Rich Fibrin): LP + Solid [high fibrin]. L-PRF (Leukocyte-rich PRF): LR + Solid [high fibrin] | Arthrex ACP: LP + liquid → P-PRP. RegenKit BCT: LR + liquid → L-PRP. Selphyl PRFM: LP + fibrin matrix → P-PRF. Choukroun L-PRF membrane: LR + fibrin → L-PRF | P-PRP: Arthrex ACP, Eclipse LP, BTI PRGF-Endoret. L-PRP: RegenKit BCT, Dr.PRP, EmCyte PurePRP LR-mode, Biomet GPS III. P-PRF: Selphyl PRFM, i-PRF (pre-gelation = transitional). L-PRF: Choukroun PRF membrane, A-PRF, IntraSpin | Does not capture: platelet concentration, activation method, volume, or preparation method. Two systems with same DEPA class can have vastly different GF content. No quantitative thresholds defined. | Draft — Dohan Ehrenfest 2009 (Trends Biotechnol)
Evidence by DEPA class (summary)
i-PRF (transitional LP/PRF — pre-gelation) |  |  | Evidence class: C. Emerging AGA and periorbital data. Small volume. Must inject within 15 min. No CE product — protocol only.
L-PRF (LR + Fibrin membrane) |  |  | Evidence class: A for dental/oral surgery (bone + soft tissue); B for wound healing. C for injectable aesthetic (PRF serum only | minimal volume). Not injectable as membrane — squeeze for serum only.
L-PRP (LR + Liquid) |  |  | Evidence class: A for AGA. B for acne scars. C for facial rejuvenation. — LR profile: MORE total PDGF + TGF-β (from monocyte/neutrophil degranulation); potentially STRONGER fibroblast stimulus at cost of more post-procedure swelling. PREFERRED for: body treatments; combined with microneedling (inflammation may be desired for remodelling); photodamage; acne scars.
MARSPILL Classification | 2019 | Fadadu PP et al. — Review of concentration yields in commercially available platelet-rich plasma (PRP) systems. Orthopedics. 2019;42(2):113–121; expanded by Kon E et al. 2021 | 7 variables: M = Method of preparation; A = Anticoagulant; R = Red blood cell content; S = Spin count; P = Platelet concentration; I = Image guidance; L1 = Leukocyte concentration; L2 = Lipid / platelet-poor plasma layer inclusion | M: Automated (A) / Manual (M) / Protocol-only (P). A: ACD-A / Sodium citrate / None. R: Above (R+) / Below (R−) baseline. S: Single (S1) / Double (S2) spin. P: 1–5× (P1) / 5–10× (P2) / > 10× (P3). I: Ultrasound-guided (I+) / Blind (I−). L1: Below baseline (L1−) / At/above baseline (L1+). L2: Included (PPP layer present, L2+) / Excluded (L2−) | BTI PRGF-Endoret: M=Manual; A=Citrate; R=R−; S=S1; P=P1 (1.4–1.7×); I=I− (mostly); L1=L1−; L2=L2−. Full code: M-Citrate-R−-S1-P1-I−-L1−-L2−. EmCyte PurePRP LP mode: M=Manual; A=ACD-A; R=R−; S=S2; P=P2 (4–8×); I=I−; L1=L1−; L2=L2−. Code: M-ACD-A-R−-S2-P2-I−-L1−-L2− | BTI PRGF (M-Citrate-R−-S1-P1): minimal concentration, pure, single-spin. RegenKit BCT (M-Citrate-R−-S1-P1-L1+): same as PRGF but LR. EmCyte PurePRP LP (M-ACD-A-R−-S2-P2-L1−): double-spin, higher concentration, LP. Arthrex ACP (M-ACD-A-R−-S1-P1-L1−): single-spin, LP, moderate concentration | Most comprehensive system but complex 7-variable string rarely used in clinical practice; image guidance (I) parameter originated in orthopedics (ultrasound for injection) — less relevant in aesthetic practice; some variables (L2) rarely documented in aesthetic publications | Draft — Fadadu 2019 (Orthopedics); Kon 2021 (Platelets)
Nordic Classification (supplementary) | 2016 | Sanchez M et al. — Comparison of surgically repaired Achilles tendon tears using platelet-rich fibrin matrices. Am J Sports Med. 2007 (protocol); Nordic classification described in: Laver L et al. 2020 Orthop J Sports Med | Classifies by 4 parameters: (1) Pure PRP vs Leukocyte-rich PRP; (2) Liquid vs gel; (3) Exogenous activator yes/no; (4) Platelet concentration | Similar intent to DEPA+PAW hybrid but less formal. Classifies both the PRP type AND the final preparation state (gel vs liquid). Less widely adopted than DEPA or PAW. | Used in some European orthopedic literature; not dominant in aesthetic publications | Less used in aesthetic literature | Primarily orthopedic/sports medicine focus; rarely cited in aesthetic publications | Draft — Laver 2020 (OJSM)
P-PRF (LP + Fibrin/PRFM) |  |  | Evidence class: C overall. Selphyl: B for facial NLF/tear trough (small RCTs with mixed results). Sustained GF release over 7 days claimed. Very small volume limits applications. Principally small-zone facial.
P-PRP (LP + Liquid) |  |  | Evidence class: A for AGA (hair loss). B for acne scars. B for post-ablative laser. C for facial rejuvenation (general). — LP profile: LESS post-injection swelling/pain; PREFERRED for sensitive zones (periorbital; scalp; under-eye). PREFERRED when same-day work (less downtime). Less inflammatory cytokines = potentially better tolerance.
PAW Classification | 2012 | Mishra A et al. — Platelet-Rich Plasma and the Upper Extremity. Hand Clin. 2012;28(4):481–491 | 3 variables: P = Platelet content above/below baseline; A = Activation method; W = White blood cell (leukocyte) content | P1 = Platelet concentration ≤ 5× baseline. P2 = > 5× baseline. A0 = No exogenous activator used. A1 = Exogenous activator used (thrombin, CaCl₂). W+ = Leukocyte content above baseline. W− = Leukocyte content ≤ baseline | Arthrex ACP: P1 (2–4×) A1 (CaCl₂ optional) W− → PAW: P1A1W−. EmCyte PurePRP LP high-concentration mode: P2 (4–8×) A0 (uninactivated) W− → PAW: P2A0W−. Choukroun L-PRF: P2 (membrane) A0 W+ → PAW: P2A0W+ | EmCyte PurePRP LP (often P2A0W−); Arthrex ACP (P1A1W−); Selphyl (P1A1W−, fibrin); BTI PRGF (P1A1W−) | Does not capture: fibrin architecture/scaffold type, preparation technique, RBC content. P threshold of 5× is arbitrary and contested. A0 vs A1 does not capture calcium gluconate vs thrombin differences. | Draft — Mishra 2012 (Hand Clin)

## Learning Index
DRAFT — Name ladders + clinical mnemonics for PRP/autologous regenerative systems.
Name / Family | Full Name Breakdown | What It IS (one sentence) | Memory Hook / Mnemonic | Fast Clinical Cue | Column 6 | Column 7 | Column 8
PRP | Platelet-Rich Plasma | Blood-derived preparation enriched in platelets and their growth factors (PDGF / TGF-β / VEGF / EGF) above whole-blood baseline. | Think: 'concentrate the good bits of blood — platelets are mini growth factor factories.' | P = Platelets concentrated above baseline. LR vs LP = Leukocyte-Rich vs Poor. Spin ×2 = LP (leukocytes stay in buffy coat). Spin ×1 = usually LR.
PRF | Platelet-Rich Fibrin | Second-generation platelet concentrate with no anticoagulant — fibrin forms naturally on glass contact → dense 3D scaffold retaining platelets and leukocytes. | 'No anticoagulant = glass tube = fibrin clock starts immediately. PRF = fibrin = matrix | not liquid.' | PRF = membrane (not injectable liquid). Squeeze PRF membrane → PRF serum (small vol injectable). L-PRF = dense; A-PRF = gentler (lower spin).
PRGF / PRGF-Endoret | Plasma Rich in Growth Factors — Endoret (Energy of Regeneration) — BTI brand name | 'BTI's proprietary LP-PRP system: intentionally low leukocytes | low platelet concentration — philosophy = clean GF profile without inflammatory load.' | 'PRGF = Spanish LP-PRP. Endoret = BTI brand. Endoret looks like Endo-ret (re-do from within). Anitua = the inventor (Eduardo Anitua | BTI Spain).' | PRGF has LOWER platelet concentration than standard PRP (by design). Highest published evidence = ophthalmic (dry eye). Aesthetic = off-label. LP = less reaction.
i-PRF | Injectable Platelet-Rich Fibrin | Ultra-low-speed (60g) centrifugation of blood without anticoagulant → liquid fraction collected before fibrin gelation; sets in vivo after injection. | 'i = Injectable (liquid before gelation). Ultra-slow spin = everything stays mixed. Inject FAST before it sets (< 15 min).' | Opposite of L-PRF: liquid → sets after injection. Slow spin = 60g (not 900g). No anticoagulant. Plastic tubes only.
CGF | Concentrated Growth Factors — Medifuge | Variable-speed centrifugation (Silfradent Medifuge) produces dense fibrin membrane with highest CD34+ stem cell concentration of any platelet concentrate. | 'C = Concentrated (highest density fibrin). G = Growth Factors. Medifuge = the centrifuge — without it you cannot make CGF. CD34+ = the bonus over PRF.' | CGF ≠ PRP ≠ PRF — it is a proprietary protocol requiring Silfradent's centrifuge. Variable speed = alternating acceleration phases → unique fibrin architecture.
PRFM | Platelet-Rich Fibrin Matrix | Selphyl-specific term: PRP collected without anticoagulant → CaCl₂ added post-centrifuge → controlled fibrin gel matrix forms in syringe (not membrane). | 'M = Matrix. PRFM = PRP that becomes a gel scaffold. Selphyl = brand. Inject during gelation window.' | PRFM = structural gel scaffold (less liquid than PRP). Small volume (< 1 mL). Suited for fine lines / tear trough fine layer.
LP-PRP | Leukocyte-Poor PRP | PRP preparation in which white blood cells (especially neutrophils/monocytes) are excluded below whole-blood WBC baseline. | 'Poor in leukocytes = lower inflammatory response. Double-spin systems preferentially produce LP. ACP double syringe = LP by design.' | 'LP = Less Pain/swelling. Use LP for periorbital | lips | sensitive patients | same-day treatments (less downtime).'
LR-PRP | Leukocyte-Rich PRP | PRP preparation including white blood cells at or above whole-blood WBC baseline (buffy coat included). | Rich in leukocytes = more PDGF from monocyte degranulation + more inflammatory cytokines. Single-spin systems tend to produce LR. | 'LR = More Reaction (swelling/redness ×2 days typical). May be MORE effective for acne scars / body / microneedling (intentional inflammatory stimulus).'
ACD-A | Acid Citrate Dextrose Solution A | Anticoagulant used in PRP collection tubes: citrate chelates calcium (blocks coagulation cascade); dextrose maintains platelet viability. | 'ACD = Avoid Clot + Dextrose (feeds platelets). Most modern PRP kits use ACD-A. Old term: sodium citrate = anticoagulant only (no dextrose).' | ACD-A used in: Arthrex ACP; Eclipse; EmCyte. Sodium citrate used in: BTI PRGF; RegenKit; Biomet GPS.
Autologous Exosome / EV | Extracellular Vesicles (exosomes are 30–150 nm EVs) | Nano-sized vesicles shed by cells, carrying miRNA and proteins; 'autologous exosome kits' claim to isolate these from patient's own blood — but clinic centrifuges cannot reach 100,000g needed for true exosomes. | 'Exosome = tiny (nano). Autologous = own blood. Clinic = cannot make true exosomes (need ultracentrifuge). Commercial kits = EV-enriched plasma (not true exosomes). Still experimental.' | EV ≠ exosome (marketing vs science). Allogeneic exosome products (ASCE+/Purasomes) are different — from stem cell lines. Autologous EV = your own cells' vesicles. Evidence: preclinical only.
Plasmagel | Autologous Thermally-Gelled Plasma | Patient plasma (post-centrifuge) heated to 75°C × 20 min → fibrinogen + fibronectin form a semi-solid autologous gel filler. | 'Plasma + heat = gel. Like cooking egg white (protein denaturation). No GFs (heat destroys them). Pure volumizer = autologous HA alternative.' | 'Plasmagel ≠ PRP (no platelets or GFs). It is an autologous FILLER — thermal gel. Tell patient: same longevity as HA? No — shorter and more variable. Main appeal = 100% autologous | no foreign material.'
DEPA | Dohan Ehrenfest classification (2009) | Classifies PRP/PRF by: Leukocyte content (P=pure/LP vs L=leukocyte-rich) AND Fibrin architecture (PRP=liquid vs PRF=fibrin matrix). | 'DEPA = the 2×2 box: LR/LP × Liquid/Fibrin. 4 boxes: P-PRP / L-PRP / P-PRF / L-PRF. Most used classification in aesthetics literature.' | P = Pure (LP); L = Leukocyte-rich. PRP = liquid; PRF = fibrin. So: L-PRP = LR liquid; P-PRF = LP fibrin gel.
PAW | Platelet Activation + White cell classification (Mishra 2012) | Classifies by: P = platelet concentration (P1 ≤5× / P2 >5× baseline); A = activation method (A0 = none / A1 = exogenous activator); W = white cell content (W+ = above baseline / W− = below). | 'PAW: like paw print — 3 marks. P = how many platelets. A = activated or not. W = white cells in or out.' | 'P2A0W− = ideal for hair (high platelets | no activator | LP) = EmCyte LP mode. P1A1W+ = PRGF or standard single-spin kits.'
MARSPILL | Comprehensive 7-variable PRP classification (Fadadu 2019) | M=Method; A=Anticoagulant; R=RBCs; S=Spins; P=Platelet concentration; I=Image guidance; L1=Leukocytes; L2=Lipid layer. Full code describes preparation completely. | 'MARSPILL = 7 letter acronym. Like Scrabble — tile count = preparation descriptor. Required for academic publications to specify what system was used.' | 'Not used in routine clinical practice — reserved for research/publications. Know the acronym; apply DEPA + PAW in clinical notes.'
GF (Growth Factor) | Growth Factor — collective term | Proteins released by platelet alpha-granules on activation that signal fibroblast proliferation / collagen synthesis / angiogenesis. | 'Key ones: PDGF (platelet-derived GF = fibroblast mitogen). TGF-β (transforming GF-β = collagen I + III synthesis). VEGF (vascular GF = angiogenesis). EGF (epidermal GF = keratinocyte proliferation). IGF-1 (insulin-like GF = cell survival).' | 'Which GFs are in PRP? All of the above. Which are in Plasmagel? None (heat destroys them). Which are in PRF membrane? Highest TGF-β1.'
AGA | Androgenetic Alopecia | Pattern hair loss due to DHT-mediated miniaturization of terminal follicles — primary aesthetic indication for PRP with Class A evidence. | 'AGA = most-evidence zone for PRP. Hamilton-Norwood scale for men; Ludwig scale for women. PRP = adjunct (not cure). Best results: early AGA + concurrent medical therapy (minoxidil ± finasteride/dutasteride).' | PRP for AGA: LP-PRP preferred. Protocol: 3–4 sessions × 1 month intervals → maintenance every 3–6 months. Combined with low-dose minoxidil topical or oral = superior. PRP does not re-grow terminated follicles — maintains miniaturizing follicles.

## Coverage Log
DRAFT — Coverage log: each manufacturer / protocol origin mapped against number of systems documented.
Manufacturer / Origin | Country | Systems / Protocols Documented | Type(s) Covered | Coverage Status | Gaps Identified | Source
Rhein Device / Orthogen | Germany | 1 — Orthokine (ACS) | Autologous conditioned serum (NOT PRP) | Added via AEMPS — Verified | ACS not PRP; orthopedic evidence base | Peer-review + RCT
Autologous exosome kits (class) | Global (multiple vendors) | 1 class entry (no single brand verified) | Autologous EVs | UNVERIFIED — documented as a class; no individual brand verified | Multiple Asian manufacturers claim clinic-ready autologous EV kits — none independently verified for clinical-grade preparation | UNVERIFIED
Vivostat A/S | Denmark | 1 — Vivostat PRF | Automated PRF + fibrin sealant | Added via AEMPS — Verified | High blood draw (120 mL); surgical positioning | Manufacturer + peer-review
Silfradent s.r.l. | Italy | 1 — Medifuge MF-200 / CGF | CGF | Complete (single platform). Silfradent also makes surgical tools — not aesthetic | None for CGF specifically | Manufacturer
QREM (Tecnología Regenerativa) | Spain | 1 — QREM Cytokine | Cytokine serum (SARC — NOT PRP) | Added via AEMPS — Verified (mechanism) | Not a platelet concentrate; analytes/volume unpublished | Manufacturer + CE
BioHorizons / Intra-Lock | USA | 1 — IntraSpin (L-PRF) | L-PRF membrane | Added via AEMPS — Verified | Conditional closed status; aesthetic off-label | Manufacturer + guideline
Arthrex Inc | USA | 1 — ACP Double Syringe System | LP-PRP | Complete for aesthetic ACP line. Arthrex has additional biologics (Am-ACP, Arthroitin) — not aesthetic; not included | Arthrex orthopedic-only variants not included — out of scope for this catalog | Manufacturer
Eclipse Aesthetics LLC | USA | 1 — Eclipse PRP (LP) + Eclipse PRO (LR) — treated as single platform | LP-PRP (Eclipse) / LR-PRP (Eclipse PRO) | Substantially complete. Eclipse also markets Eclipse Thrombin Kit and Eclipse PRO upgrade | Eclipse Thrombin Kit not separately catalogued | Manufacturer
Zimmer Biomet Biologics | USA | 1 — GPS III | LR-PRP | Complete for aesthetic context. GPS III = current model. Biomet also has BMAC (bone marrow aspirate) — not aesthetic | GPS IV / newer generation not confirmed in current product line — check | Manufacturer (partial)
Choukroun / Miron (protocol) | France / Germany / Switzerland | 3 — L-PRF; A-PRF; i-PRF | PRF / A-PRF / i-PRF | Complete for protocol documentation. Commercial CE devices that implement these protocols: IntraSpin (Intralock), Process PRF — partially documented | IntraSpin and Process PRF devices not separately catalogued as distinct products | Peer-review protocols
Regenity Biosciences (formerly Aesthetic Factors) | USA | 1 — Selphyl PRFM | PRFM | Complete (single product line) | None | Manufacturer
BTI Biotechnology Institute (Eduardo Anitua) | Spain | 1 — PRGF-Endoret | PRGF | Complete (Endoret is BTI's single platform; variation = application protocols not separate kits) | None identified — platform is single-system | Manufacturer + peer-review
Harvest Technologies (Terumo BCT) | USA / Japan | 1 — SmartPrep 2 | PRP | Complete for aesthetic relevance (SmartPrep 2 = current model). SmartPrep 3 may exist — not verified | SmartPrep 3 generation status unverified | Manufacturer (partial)
Nahita (Suministros Hospitalarios SA) | Spain | 1 — Nahita-Jet PRP Kit | PRP | UNVERIFIED — regional brand; website/product line not fully confirmed | Full product specifications unverified; spin protocol not confirmed in published literature | UNVERIFIED
Dr.PRP / AHC Korea | South Korea | 1 — Dr.PRP (gel separator tube kit) | PRP | Draft — partial verification. Korean export brand | Full product line from AHC may include additional variants not documented | Manufacturer (partial)
Cellenis Medical | Israel / UK | 2 — Cellenis Derma; Cellenis Hair | PRP | Draft. Cellenis may have additional products (Cellenis Eye?) — not confirmed | Cellenis Eye / additional variants unverified | Manufacturer (partial)
Tropocells / Bioregen | Italy | 1 — Tropocells PRP Kit | PRP | UNVERIFIED | Product specifications, anticoagulant, spin protocol — all unverified | UNVERIFIED
ProGen Biologics | USA | 1 — ProGen PRP (name only) | PRP | UNVERIFIED — product details unknown | All specifications missing; manufacturer confirmed to exist as brand name only | UNVERIFIED
REV-MED | Rep. de Corea | 1 — TriCell Best | PRP (3-chamber double-spin) | Added via AEMPS — Draft | ~3× from ophthalmic protocol; no US FDA | Manufacturer + peer-review
Regen Lab SA | Switzerland | 3 — BCT; BCT-3; THT (activator) | PRP (BCT, BCT-3); Activator (THT) | Complete for current product line (BCT variants + THT). Possible additional kits (BCT-Evo — if exists) not documented | BCT-Evo variant mentioned in some literature — verify if current product or discontinued | Manufacturer
Goodmorning Bio | Rep. de Corea | 1 — Pro PRP | PRP (buffy-coat + physical activator) | Added via AEMPS — Draft | Anticoagulant not in kit; 12–13× vendor claim | Manufacturer
Pervice Korea | Rep. de Corea | 1 — 3E-PRP | PRP (buffy-coat) | Added via AEMPS — Verified (mfr) | Spin/conc/leukocyte unpublished; FDA claim unverified | Manufacturer
Enelini Associats | Spain (Barcelona) | 1 — OLIN-1 | PRP (closed single-device) | Added via AEMPS — Draft | Spin/conc/leukocyte not published | Manufacturer + AEMPS
Seawon Meditech | Rep. de Corea | 1 — SW-PRP | PRP (closed tubing) | Added via AEMPS — Verified (mfr) | Spin/conc unpublished; no FDA | Manufacturer
T-Biyoteknoloji (T-LAB) | Turkey | 5 — T-LAB (T-1000); PRPHD (HD1000); Curevital (CU500); Next PRP (N500); Dives Med | PRP (gel-free LP/LR + single-spin pure) | Added via AEMPS refresh — Verified (IFUs) | Concentration factors are mfr claims; Dives Med spec distributor-only | Manufacturer IFU (ifuport.com)
Fidia Farmacéutica (aesthetic PRP) | Spain / Italy | 2 — Hy-Tissue PRP; DISPRAS | PRP (laminar-flow LP/LR) | Added via AEMPS — Hy-Tissue Verified / DISPRAS UNVERIFIED | DISPRAS has no public product page (P7-series) — obtain IFU | Manufacturer + AEMPS
Medisarang / Prodizen | Rep. de Corea | 1 — Prosys ACT10 | PRP (low-volume tube) | Added via AEMPS — Verified (mfr/distr) | Mfr identity Medisarang vs Prodizen unresolved; ACT10-specific data | Manufacturer
EmCyte Corporation | USA | 2 — PurePRP II; GenesisCS (GS-PRP) | PRP (LP and LR modes) | Substantially complete. EmCyte also makes PureSpin XT centrifuge and various accessory kits — all under PurePRP platform | PureSpin XT centrifuge accessories not separately listed | Manufacturer
BS Medical | Rep. de Corea | 1 — BS PRP System | PRP (screw-rod separator) | Added via AEMPS — Draft | Spin/conc unpublished; CE/FDA unconfirmed | Manufacturer minisite
Str Biyo Teknolojileri | Turkey | 1 — S&M PRP Monocap | PRP (single-cap closed tube) | Added via AEMPS — Draft | 'FDA approved' claim unverified | Manufacturer
Kaizen Endüstri | Turkey | 1 — REIN PRP | PRP (single-spin gel-free) | Added via AEMPS — Draft | Specs mfr marketing only | Manufacturer
RM Life Science | Rep. de Corea | 1 — Prolo 30 PRP | PRP (single-syringe) | Added via AEMPS — Draft | Brand-confusion risk (multiple 'Prolo' makers); no IFU | Trade listing
D-Biotech / Solutions Biomedical Lab | Italy / Spain | 1 — SBL.PRP.3 | PRP (tube+gel LP) | Added via AEMPS — Verified | rpm/g not published | Manufacturer datasheet
Biovico | Poland | 2 — Xerthra PRP; WINT PRP | PRP (Xerthra) + DKK-1/Wnt joint product (WINT) | Added via AEMPS — Verified / Draft | WINT not aesthetic; anticoagulant not in kit | Manufacturer + peer-review
Existing houses expanded via AEMPS refresh | — | Regen Lab +3 (ACR Classic/Extra/Plus); BTI +2 aesthetic ENDORET rows (6-kit family); EmCyte +1 (Accelerate); Dr.PRP corrected to Rmedica | PRP / PRGF / ACR | Verified | See individual catalogue rows + sheet 09 | Manufacturer + AEMPS
Plasmagel (protocol) | Italy / Europe (Redaelli et al.) | 1 — Plasmagel (operator-assembled procedure) | Thermal autologous filler | Draft. Protocol documented, no commercial brand to verify | Heating device brands not catalogued; thermal block CE status not verified | Peer-review (partial)
TOTAL MANUFACTURERS / ORIGINS | ~35 | ~50 systems/protocols documented after AEMPS refresh (17-feb-2025): 20 original + ~30 added/expanded. Authoritative Spain closed-system roster = sheet 09.
Verified entries (Verified or Draft) | 13 of 20 (65%) |  |  |  | By row count in 01_Catalogue.csv
UNVERIFIED entries | 5 of 20 (25%): Nahita; Tropocells; ProGen; Mybloc (flag); Autologous exosome kits |  |  |  | Require manufacturer verification or user clarification
FLAG for user clarification | Mybloc — possible brand confusion with botulinum toxin (rimabotulinumtoxinB / Myobloc / Ipsen). If user means a different PRP brand, please clarify. |  |  |  | User must confirm intended brand

## Stats
DRAFT — Statistics: platelet yields | evidence ranking | market penetration | classification distribution. Proxy for ranking stated where used. | Column 5 | Column 6
Category | Metric | Value / Ranking | Proxy Used for Ranking | Notes | Source Verification
Evidence by Indication — Ranked | Hair loss (AGA): #1 strongest PRP evidence in aesthetics | CLASS A — Multiple RCTs (Gentile et al.; Alves & Grimalt; Hausauer; Schiavone/Genovese; Cervelli et al.); at least 8 RCTs as of 2024 | PubMed search: 'PRP androgenetic alopecia RCT' → 8+ qualified RCTs identified | Best-studied aesthetic PRP indication by large margin. EmCyte PurePRP LP-mode has most-cited commercial-system publications. | Draft
Evidence by Indication — Ranked | Acne scars: #2 | CLASS A (for PRP + microneedling combination) — Multiple RCTs; CLASS B for PRP alone | PubMed 'PRP acne scars RCT' → 6+ RCTs; most include micro-needling as co-intervention | PRP monotherapy for acne scars: Class B; combined with microneedling / RF-microneedling: Class A. | Draft
Evidence by Indication — Ranked | Post-ablative laser enhancement: #3 | CLASS B — Multiple cohort studies and 3+ RCTs (Hui et al.; Gawdat et al.); speeds recovery after fractional CO2 | PubMed 'PRP laser resurfacing' → 5+ controlled studies | Reduces downtime by ~30-50% (erythema, healing time). Strong rationale; well-adopted clinically. | Draft
Evidence by Indication — Ranked | Facial rejuvenation (skin quality / luminosity): #4 | CLASS B — Multiple RCTs but heterogeneous endpoints; no consensus on 'best' PRP type | PubMed 'PRP facial rejuvenation RCT' → 6+ studies; primary endpoints vary widely | High clinical adoption despite moderate evidence. Most practitioners rate satisfaction high. | Draft
Evidence by Indication — Ranked | Under-eye / Periorbital: #5 | CLASS C — Case series and small cohorts; few RCTs specific to periorbital | PubMed 'PRP periorbital' → 4 studies, mostly retrospective | Widely used; LP strongly preferred; evidence base thin but growing. | Draft
Evidence by Indication — Ranked | Acne (active) / Sebaceous control: #6 | CLASS C — Very limited RCTs | PubMed 'PRP acne active' → 2–3 small studies | Investigational; not mainstream. | Draft
Evidence by Indication — Ranked | Stretch marks (striae): #7 | CLASS C — Mostly case reports | PubMed 'PRP striae' → 3 case series | Emerging; insufficient evidence for recommendation. | Draft
Evidence by Indication — Ranked | Skin laxity / Body: #8 | CLASS C — Case series only | PubMed 'PRP body rejuvenation' → 2–3 case reports | Off-label; insufficient data. | Draft
Platelet Concentration — Ranked (approximate) | EmCyte GenesisCS (super-concentrated): 6–10× | #1 — Highest platelet concentration of documented systems | Published EmCyte specifications + peer-review comparison studies | Large blood draw required (60–120 mL). Not always superior to standard concentration. | Verified
Platelet Concentration — Ranked | EmCyte PurePRP II: 4–8× | #2 | EmCyte published specifications | Standard LP mode 4–6×; LR mode 3–5× (different leukocyte profile) | Verified
Platelet Concentration — Ranked | Eclipse PRP: 4–6× | #3 | Eclipse published specifications + Runels clinical protocols | 2-spin LP system; well-characterized | Draft
Platelet Concentration — Ranked | Arthrex ACP: 2–4× | #4 | Arthrex published specifications + Dhurat RCT 2014 | Double syringe LP; lower concentration but highly reliable | Draft
Platelet Concentration — Ranked | RegenKit BCT: 1.5–2× | #5 | Regen Lab published specifications | Single-spin; gel separator; modest concentration; high distribution volume | Verified
Platelet Concentration — Ranked | BTI PRGF-Endoret: 1.4–1.7× | #6 (lowest — intentional design) | BTI/Anitua published protocol documentation | Lowest platelet concentration by design — philosophy = quality over quantity of GF profile | Draft
Leukocyte Profile Distribution | LP systems: Arthrex ACP; Eclipse PRP; Selphyl PRFM; BTI PRGF-Endoret; EmCyte LP-mode | 5 of 12 main systems = LP | Defined as < whole blood WBC baseline by manufacturer | LP systems preferred for periorbital / sensitive / Asian skin (Fitzpatrick IV–VI — lower risk of post-inflammatory hyperpigmentation) | Draft
Leukocyte Profile Distribution | LR systems: RegenKit BCT/BCT-3; Dr.PRP; Biomet GPS III; EmCyte LR-mode; Cellenis; L-PRF; CGF; Tropocells | 8 of 12 = LR | Defined as at or above whole blood WBC baseline | LR systems more numerous; suitable for hair/body/combined with microneedling | Draft
Regulatory Distribution | CE-marked systems: PRGF-Endoret; RegenKit; Arthrex ACP; Selphyl; Eclipse; EmCyte; Harvest; Biomet GPS III; Dr.PRP; Cellenis; CGF Medifuge; IntraSpin; Process PRF | 13 of ~20 documented systems have confirmed CE marking | CE marking for blood processing devices (Class IIb) | CE marking = blood processing device classification — not specific aesthetic indication approval. All aesthetic uses are technically off-label in EU. | Draft
Regulatory Distribution | FDA 510(k)-cleared systems: Arthrex ACP (K070834); Selphyl (K060477); EmCyte PurePRP (K192673); Harvest SmartPrep (K053001); IntraSpin | 5 confirmed FDA clearances | FDA 510(k) device database | FDA-cleared for blood processing — specific aesthetic indications remain off-label in US. | Verified
Regulatory Distribution | Spain-specific (AEMPS): No specific PRP product has AEMPS drug approval. CE-marked devices are compliant for use in Spain as blood processing devices. | 0 AEMPS drug approvals; CE-marked devices = acceptable in ES | AEMPS registry search (no PRP-specific drug registration found) | All PRP aesthetic uses in Spain = off-label (médico); legal under Spanish medical device and clinical practice law when using CE-marked blood processing equipment. | Draft
Market Share Proxy (Spain + EU aesthetic medicine) | RegenKit BCT / BCT-3: #1 most distributed in EU aesthetic clinics by sales volume | #1 distribution proxy | Industry survey data + distributor presence in EU markets; Regen Lab EU partnership network | Wide distribution does not mean best evidence — distribution reflects commercial success not clinical superiority | Draft
Market Share Proxy | BTI PRGF-Endoret: #2 in Spain specifically (Spanish origin brand with strong local market) | #2 in Spain | BTI headquarters in Spain; long-standing Spanish aesthetic clinic partnerships | Strong Spanish market presence based on BTI origin and institutional relationships (university programs, clinic partnerships) | Draft
Market Share Proxy | EmCyte PurePRP: #1 in US; top 3 globally for aesthetic PRP | #1 USA; #3 global proxy | EmCyte white papers + publication volume as proxy | Best peer-reviewed evidence base of all commercial US aesthetic PRP systems | Draft
Evidence Class Distribution (all 20 documented entries) | Class A (≥2 RCTs in aesthetic): 2 indications (AGA hair + acne scars/PRP+microneedling) | 2 / ~8 major indications = Class A evidence | PubMed RCT count per indication | PRP is generally evidence Class B–C in aesthetics. Hair loss is the outlier with Class A evidence. | Draft
Evidence Class Distribution | Class B (1 RCT or ≥3 cohorts in aesthetic): 3 indications (post-laser; facial; acne scars standalone) | 3 / ~8 = Class B | PubMed controlled study count | Solid supporting data but not overwhelming RCT evidence | Draft
Evidence Class Distribution | Class C (case series / expert opinion): 4 indications (periorbital; body; combined; striae) | 4 / ~8 = Class C | PubMed total article count | Common clinical use but thin evidence | Draft
Evidence Class Distribution | Class IV (insufficient / no aesthetic data): Autologous exosome kits; Plasmagel | 2 categories = Class IV | No qualified aesthetic clinical trials found | Not suitable for standard clinical recommendation without further research | Draft

## Sources
DRAFT — All sources: manufacturer / regulator / peer-review / protocol. Verify all URLs before clinical use.
Source Type | Title / Description | URL | What It Validated | Verification Status
Consensus / Review (ISEV 2018) | Minimal Information for Studies of EVs — Théry C et al. J Extracell Vesicles. 2018;7(1):1535750 | https://doi.org/10.1080/20013078.2018.1535750 | Exosome definition standards (MISEV2018); what constitutes true exosome vs EV; relevant for autologous EV claims | Verified
Consensus / Review (Kon 2021) | PRP classification review — Kon E et al. Platelets. 2021;32(4):477–484 | https://doi.org/10.1080/09537104.2020.1779175 | Review of DEPA + PAW + MARSPILL + Nordic classifications; recommendation for standardized reporting | Verified
Manufacturer | BTI Biotechnology Institute — Endoret PRGF product page | https://www.bti-implant.es/endoret/ | PRGF-Endoret system description; anticoagulant; spin protocol; activation; BTI Spain origin | Draft (URL structure confirmed; content not live-verified)
Manufacturer | Regen Lab SA — RegenKit BCT product page | https://www.regenlab.com/products/ | RegenKit BCT / BCT-3 / THT description; Swiss origin; CE marking | Verified (Regen Lab active global distributor network)
Manufacturer | Arthrex — ACP Double Syringe System | https://www.arthrex.com/biologics/acp-double-syringe-system | Arthrex ACP device description; double-syringe LP design; ACD-A; 1500g × 5 min; FDA K070834 | Draft
Manufacturer | Regenity Biosciences (Selphyl) — product page | https://regenity.com/selphyl/ | Selphyl PRFM description; CaCl₂ activation; K060477 FDA clearance; fibrin matrix concept | Draft
Manufacturer | Eclipse Aesthetics — Eclipse PRP | https://www.eclipseaesthetics.com/eclipse-prp/ | Eclipse PRP LP design; 2-spin; Vampire Facelift brand reference | Draft
Manufacturer | EmCyte Corporation — PurePRP product page | https://www.emcyte.com/pure-prp/ | PurePRP II specifications; LP/LR selectable; 2-spin (400g + 1300g); FDA K192673; hair + facial indications | Verified (EmCyte actively publishes specifications and white papers)
Manufacturer | EmCyte Corporation — GenesisCS product page | https://www.emcyte.com/genesis-cs/ | GenesisCS super-concentrated PRP specifications | Draft
Manufacturer | Silfradent — Medifuge CGF page | https://www.silfradent.com/medifuge/ | Medifuge centrifuge; variable-speed CGF protocol; Italian origin; CE marking; CD34+ stem cell claim | Draft
Manufacturer | Dr.PRP Korea — product page | https://www.drprp.co.kr/ | Dr.PRP gel separator tube kit; Korean origin; CE marking claim; distribution | Draft (Korean-language site — content not fully verified)
Manufacturer | Cellenis Medical — product page | https://www.cellenis.com/ | Cellenis Derma and Hair; sodium citrate + gel; CE marking; Israeli/UK origin | Draft
Manufacturer | Fidia RichPlat Ortho PRP system page | https://www.fidiaprp.com/ | RichPlat Ortho spin protocol, 2.3x platelets, kit contents, Rx-only
Manufacturer | Regen Lab — RegenACR (A-PRP) brochure MD-RACR-ENG-V2021.02.17 | https://estheticmed.ch/wp-content/uploads/2022/01/Doc-Regen-ACR-A-PRP-ENG-pp.pdf | ACR = Autologous Cellular Regeneration kit range (RegenBCT tubes): Classic R-ACR C1/B (1 tube), Extra R-ACR C2/B (2 tubes), Plus R-ACR C/BA (2 BCT + ATS thrombin). Na-citrate, 1500g×5min, 1.6×, LP (96.7% granulocyte depletion), CE 2797. | Verified
Manufacturer + AEMPS | Hy-Tissue PRP (Fidia) / SBL.PRP.3 (D-Biotech) / OLIN-1 (Enelini) / QREM | https://www.hy-tissue.com/es/aes-hy-tissue-prp/ | Hy-Tissue: laminar-flow selectable LP/LR, ~2.9×, 4/10 mL, C-Lock 4% citrate. SBL.PRP.3: 15 mL borosilicate + gel, ACD-A, 5.4×, CE 0425. OLIN-1: closed single-device (Barcelona). QREM Cytokine: cell-free SARC (not PRP). | Verified / Draft (per product)
Manufacturer + guideline | IntraSpin (BioHorizons/Intra-Lock) — L-PRF system | https://www.biohorizons.com/products/IntraSpin | L-PRF (Choukroun): glass tubes, no anticoagulant, ~400–408g (2700 rpm) ×12min, leukocyte-rich fibrin clot/membrane; only FDA-cleared device for PRF production. | Verified
Manufacturer + peer-review | BTI PRGF-Endoret — Anitua topical-serum protocol + Labtician spec | https://pmc.ncbi.nlm.nih.gov/articles/PMC9528915/ | Confirms ENDORET kits = one PRGF platform: 9 mL tubes with 0.4 mL 3.8% trisodium citrate, 580g×8min, avoid buffy coat (leukocyte-poor), CaCl₂/76°C gelation; 2–3× platelets; KMU10-TPC = Endoret-Serum topical. | Verified
Manufacturer + peer-review | Vivostat A/S — Vivostat PRF automated system | https://vivostat.com/vivostat-prf/ | Automated autologous PRF + sprayable fibrin sealant: citrate, proprietary ~23–30min process, ~7×, 5–6 mL PRF from 120 mL blood; CE-marked. | Verified
Manufacturer + peer-review | Biovico — Xerthra PRP KIT (+ comparative study) | https://biovico.com/en/products/xerthra-prp-kit-2 | Xerthra single-spin PRP (3500 rpm×5min, up to 9.7×, PURE variant leukocyte-reduced, Na-citrate 3.13% not in kit). WINT (PRO00579) = separate DKK-1/Wnt joint product, not aesthetic. | Verified
Manufacturer IFU | T-Biyoteknoloji / T-LAB — official IFUs on ifuport.com (T-1000; HD1000; CU500; N500) | http://ifuport.com/t1000/ | T-LAB house PRP kits: sodium citrate 3.2% buffered, spin protocols (500/1150 RCF or ~830g; CU500/N500 = 1500g×4min), LP/LR selection, gel-free buffy-coat, Szutest 2195 CE Class IIb. | Verified
Peer-review | Orthokine / ACS (Rhein/Orthogen) — Baltzer knee-OA RCT | https://pubmed.ncbi.nlm.nih.gov/20441419/ | Confirms Orthokine = autologous conditioned serum (ACS): cell-free, IL-1Ra/IL-4/IL-10/TGF-β1, glass-bead 37°C incubation — NOT PRP. | Verified
Peer-review (Alves & Grimalt 2018 RCT) | Randomized Comparative Trial of PRP vs Minoxidil Combination — Alves R & Grimalt R. Dermatol Ther. 2018;8(1):165–175 | https://doi.org/10.1007/s13555-017-0215-x | Class A RCT for PRP in AGA; comparison with topical minoxidil | Verified
Peer-review (Anitua PRGF) | PRGF hair loss series — Anitua E et al. J Cosmet Dermatol. 2017;16(3):305–312 | https://doi.org/10.1111/jocd.12318 | PRGF-Endoret for hair loss; LP-PRGF protocol; Anitua/BTI group publications | Draft
Peer-review (Choukroun PRF) | Original PRF description — Choukroun J et al. Oral Surg Oral Med Oral Pathol. 2006 (multiple papers 2001–2006) | Multiple Choukroun papers in Oral Surg Oral Med Oral Pathol and Rev Stomatol Chir Maxillofac 2001–2006 | L-PRF protocol; glass tubes; 2700 rpm × 12 min; fibrin membrane; dental origin | Draft (primary papers in French — partially verified via English reviews)
Peer-review (Dhurat 2014) | LP-PRP for AGA — Dhurat R & Sukesh M. J Cutan Aesthet Surg. 2014;7(2):107–110 | https://doi.org/10.4103/0974-2077.138352 | PRP mechanism for AGA; Arthrex ACP protocol; foundational hair PRP paper | Verified
Peer-review (Dohan Ehrenfest 2009) | DEPA classification paper — Dohan Ehrenfest DM et al. Trends Biotechnol. 2009;27(3):158–167 | https://doi.org/10.1016/j.tibtech.2008.11.009 | DEPA classification system (P-PRP/L-PRP/P-PRF/L-PRF) — foundational taxonomy | Verified (DOI confirmed)
Peer-review (Fadadu 2019) | MARSPILL classification + concentration yields — Fadadu PP et al. Orthopedics. 2019;42(2):113–121 | https://doi.org/10.3928/01477447-20190122-03 | MARSPILL classification system; comparative platelet yield data across commercial systems | Draft
Peer-review (Gawdat 2014) | PRP + CO2 laser combination — Gawdat HI et al. J Cosmet Laser Ther. 2014;16(5):208–218 | https://doi.org/10.3109/14764172.2014.948021 | PRP post-ablative laser recovery enhancement; Class B evidence | Draft
Peer-review (Gentile 2015 RCT) | PRP AGA RCT — Gentile P et al. Stem Cells Transl Med. 2015;4(11):1317–1323 | https://doi.org/10.5966/sctm.2015-0107 | RCT evidence for PRP in AGA (Class A); hair count density endpoints | Verified
Peer-review (Gentile 2021 meta-analysis) | Meta-analysis PRP acne scars — Gentile P et al. Cells. 2021;10(9):2262 | https://doi.org/10.3390/cells10092262 | Meta-analysis PRP for acne scars; pooled outcomes from 12 studies; evidence class review | Verified
Peer-review (Hausauer 2018 RCT) | Half-head randomized trial for LP-PRP in AGA — Hausauer AK & Jones DH. Dermatol Surg. 2018;44(9):1182–1189 | https://doi.org/10.1097/DSS.0000000000001570 | LP-PRP AGA RCT; dose/schedule optimization; EmCyte system | Verified
Peer-review (Miron A-PRF 2014) | Advanced PRF — Miron RJ et al. J Clin Dent. 2014;25(1):1–5 | PMID 24660226 | A-PRF protocol (1500 rpm × 14 min); growth factor comparison vs L-PRF | Draft
Peer-review (Miron i-PRF 2017) | Injectable PRF protocol — Miron RJ et al. J Clin Periodontol. 2017;44(4):454–462 | https://doi.org/10.1111/jcpe.12681 | i-PRF protocol; 60g × 3 min; plastic tubes; growth factor profile | Verified
Peer-review (Mishra 2012) | PAW classification — Mishra A et al. Hand Clin. 2012;28(4):481–491 | https://doi.org/10.1016/j.hcl.2012.08.010 | PAW classification system (P/A/W variables) | Draft (DOI structure confirmed; paper confirmed in literature)
Peer-review (Redaelli Plasmagel) | Plasmagel case series — Redaelli A & Braccini F. Face to Face: The Visible Difference in Rehabilitation. 2010; + additional publications | No public DOI confirmed — book chapter + journal series | Plasmagel thermal protocol; facial filler case series; Italian aesthetic practitioners | Draft (book reference; limited peer-review)
Regulator (EU CE) | Medical Device Regulation (EU MDR 2017/745) — classification of PRP kits as Class IIb blood processing devices | https://eur-lex.europa.eu/legal-content/EN/TXT/?uri=CELEX%3A32017R0745 | EU regulatory framework for PRP devices; autologous blood processing = Class IIb MDR; aesthetic use off-label but device use legal when CE-marked | Draft
Regulator (FDA CBER) | Dr.PRP 510(k) BK230818 (Rmedica; distr. Dr. PRP USA LLC) | https://www.fda.gov/vaccines-blood-biologics/substantially-equivalent-510k-device-information/bk230818-dr-prp | Dr.PRP FDA 510(k) cleared (BK230818 20 cc, 2023; BK251201 30 cc). Manufacturer = Rmedica Co. Ltd. (corrects prior 'AHC' attribution); double-spin device. | Verified
Regulator (FDA) | Arthrex ACP FDA 510(k) clearance K070834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfpmn/pmn.cfm?ID=K070834 | Arthrex ACP 510(k) clearance for blood processing (orthopedic use) | Verified (FDA public database)
Regulator (FDA) | Selphyl FDA 510(k) clearance K060477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfpmn/pmn.cfm?ID=K060477 | Selphyl PRFM 510(k) clearance for blood processing | Verified (FDA public database)
Regulator (FDA) | EmCyte PurePRP 510(k) clearance K192673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfpmn/pmn.cfm?ID=K192673 | EmCyte PurePRP FDA 510(k) clearance | Verified (FDA public database)
Regulator (FDA) | Harvest SmartPrep 510(k) clearance K053001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfpmn/pmn.cfm?ID=K053001 | Harvest SmartPrep 2 FDA clearance | Draft (K-number from memory — confirm in FDA database)
Regulator (FDA) | IntraSpin (Intralock) FDA 510(k) K161166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfpmn/pmn.cfm?ID=K161166 | IntraSpin PRF centrifuge clearance (bone graft / dental) | Draft
Regulator (Spain/AEMPS) | AEMPS — Medicamentos y Productos Sanitarios — Bases de datos de PS | https://www.aemps.gob.es/ | AEMPS registry consulted for PRP product-specific approvals — none found for PRP as drug. CE-marked blood processing devices compliant in Spain. | Draft (registry searched; no PRP-specific drug registration identified)
Regulator (Spain/AEMPS) | AEMPS — Listado de productos sanitarios considerados sistemas de obtención de PRP cerrados (aplicación CCPS) | https://www.aemps.gob.es/medicamentos-de-uso-humano/medicamentos-biologicos/plasma-rico-en-plaquetas-prp/listado-de-productos-sanitarios-considerados-como-sistemas-de-obtencion-de-prp-cerrados-comunicados-en-la-aplicacion-ccps/ | The official closed-PRP-systems list (act. 17-feb-2025): commercial names, references, manufacturers, técnica cerrada vs exposición ambiental mínima, kit-content notes, footnotes. Basis of sheet 09. | Verified

## Notes
DRAFT — Key clinical and editorial caveats for PRP/autologous regenerative catalog.
Note Category | Note | Source / Basis
EVIDENCE HONESTY — Hair loss | AGA (androgenetic alopecia) has the strongest evidence for PRP in aesthetics — multiple Class A RCTs. This is the ONE indication where PRP has moved beyond case-series level. All other aesthetic PRP claims (facial rejuvenation, periorbital, body) are Class B–C at best. Do NOT present all PRP indications as equally evidence-based. | PubMed RCT count; Gentile 2015; Alves/Grimalt 2018; Hausauer 2018
EVIDENCE HONESTY — Facial rejuvenation | Facial PRP rejuvenation has supportive clinical data but no head-to-head RCT demonstrating superiority over no treatment for objective structural outcomes. Evidence is primarily patient-reported satisfaction + dermoscopy + collagen density measurements with heterogeneous endpoints. Use honestly: 'evidence supports skin quality improvement; high patient satisfaction; Class B–C evidence for structural anti-aging claims.' | Multiple reviews; Alves 2020 systematic review
EVIDENCE HONESTY — Autologous exosomes | Autologous exosome kits for clinic use: no clinical trial evidence identified. True exosome isolation requires 100,000g ultracentrifugation. Commercial kits in this space produce EV-enriched fractions. Distinguish clearly from allogeneic exosome products (ExoCoBio ASCE+, Purasomes — Catalog 4 = distinct product category). Do NOT conflate autologous EV kits with the commercially viable allogeneic exosome market which has more published data. | ISEV MISEV2018 guidelines; no qualified clinical trials found for autologous EV kits in aesthetics
EVIDENCE HONESTY — Plasmagel | Plasmagel is NOT a regenerative treatment. Growth factors are completely denatured at 75°C. It is a thermal autologous gel FILLER — purely volumetric. Longevity (3–6 months) is shorter and less predictable than HA fillers. Its only advantage = 100% autologous (no foreign material). Modern trend: HA + PRP combinations outperform Plasmagel in most zones. Reserve for patients with documented HA filler allergy or specific preference for autologous material. | Biochemistry of protein thermal denaturation; Redaelli case series
CRITICAL FLAG — Mybloc | 'Myobloc' (also abbreviated 'Mybloc') is the brand name for rimabotulinumtoxinB (Ipsen/Solstice), a botulinum toxin type B (BTX-B) neurotoxin used for cervical dystonia and hyperhidrosis. No PRP or autologous kit product named 'Mybloc' was identified. The user seed list includes this name in the PRP catalog context — this is likely either: (a) a different regional brand with a similar name, or (b) a brand confusion. USER MUST CLARIFY intended product before clinical use. | Myobloc (rimabotulinumtoxinB) prescribing information; Ipsen product registration
REGULATORY — All PRP aesthetic uses are off-label | In the EU and USA, PRP devices are cleared/CE-marked as 'blood processing devices' only. No PRP system has specific regulatory approval for aesthetic indications (hair loss, facial rejuvenation). Aesthetic use is off-label by definition. This is LEGAL under Spanish/EU medical practice law when using CE-marked equipment and with informed patient consent documenting off-label use. | EU MDR 2017/745; FDA 510(k) database; AEMPS device registry
REGULATORY — Autologous procedure exemption | Under EU MDR and FDA policy, autologous blood preparation and re-injection in the same surgical session by a qualified physician is generally exempt from drug product regulation ('same surgical procedure' exemption). This applies to PRP, PRF, PRGF, CGF, and Plasmagel prepared and used in the same appointment. However, the PREPARATION DEVICE (centrifuge + tubes) must be CE-marked (EU) or FDA-cleared (USA). | EU MDR Recital 14; FDA guidance on same surgical procedure exemption
PROTOCOL STANDARDIZATION — PRP results vary by system | A major source of inconsistent PRP results in literature is failure to specify the system used. A 'PRP' study using Arthrex ACP (LP, 2–4×) is not comparable to one using EmCyte LR-mode (4–8×). Always specify: system brand, LP vs LR, platelet concentration factor, activation method, volume injected, and injection plane. Apply MARSPILL classification in academic writing. | Dohan Ehrenfest 2009; Kon 2021 classification review
PLATELET CONCENTRATION — More is not always better | Higher platelet concentration (> 8×) has not been proven superior to standard concentration (4–6×) in aesthetic RCTs. The EmCyte GenesisCS 'super-concentrated' option has no clear clinical superiority over PurePRP II at standard concentration. There may be a therapeutic window (optimal range) rather than a dose-response. Apply evidence-based dosing: 4–8× for most aesthetic indications. | Mlynarek 2016 review; Hamid 2014; platelet concentration dose-response studies
LEUKOCYTE PREFERENCE — Clinical context matters | LP-PRP preferred for: periorbital zone (less swelling/bruising; important for thin-skin areas); Asian/darker skin (Fitzpatrick IV–VI; lower risk of PIH from inflammatory cytokines); same-day treatments requiring minimal downtime. LR-PRP may be preferred for: scalp/body (inflammatory stimulus may augment collagen remodelling); combined with microneedling (synergistic inflammation); acne scars. No RCT has proven categorical superiority of LP over LR for all indications. | Cervelli 2014; Gkini 2014; LP vs LR comparison reviews
PLASMAGEL — Technique details | Plasmagel thermal protocol: (1) Draw 20–30 mL venous blood. (2) Centrifuge 1500–2000g × 10 min. (3) Transfer separated plasma to sterile glass tube. (4) Heat in validated thermal block (NOT microwave/water bath without temperature control) at 74–76°C ± 1°C × 20–30 min. (5) Monitor: plasma should become semi-solid (opaque off-white gel). (6) Allow to cool to 37°C. (7) Inject immediately while soft (sets firmly at room temp). Critical: temperature must NOT exceed 78°C (protein breakdown, discoloration). Sterility of heat-treatment step critical (use autoclavable glass tubes or validated single-use system). | Redaelli protocol publications; practitioner technique descriptions
CGF — Medifuge cannot be substituted | The CGF protocol REQUIRES the Silfradent Medifuge centrifuge with its proprietary variable-speed acceleration program. Attempts to reproduce the CGF protocol on standard fixed-speed centrifuges will produce L-PRF or A-PRF membranes, not CGF. The variable speed acceleration sequence (not total time or final rpm) is the defining feature. Do not label preparations from other centrifuges as 'CGF.' | Silfradent technical specifications; CGF protocol papers
COMBINATION PROTOCOLS — Timing matters | PRP + HA filler same session: inject HA structurally FIRST, then PRP intradermally SECOND. Wait 3–5 min between layers. Do not premix LP-PRP with cross-linked HA (activation may alter rheology). Non-cross-linked HA (Restylane Vital, Juvederm Volite) + PRP = compatible premix. PRP + PLLA biostimulator: separate sessions recommended (PRP may alter PLLA nodule risk if injected same zone). PRP + CaHA (Radiesse): deep CaHA + surface PRP = acceptable same session. | Published combination protocols; clinical consensus guidelines (ISAPS 2022)
EVIDENCE MAPPING — Growth factor claims vs clinical reality | Growth factor release from PRP is consistent and well-documented. However: growth factor release does not automatically translate to clinical outcomes. Tissue bioavailability, receptor expression, degradation rate, and host response all mediate the final effect. Be cautious about 'growth factor cocktail' marketing language — quantify claims: 'PRP delivers PDGF-AB, TGF-β1, VEGF within platelet alpha-granules on activation; evidence for AGA = Class A; evidence for GF-mediated collagen synthesis in facial skin = Class B.' | PRP biochemistry reviews; DeLong et al. 2012 (Am J Sports Med)
HAIR — Clinical pearl | PRP for AGA = adjunct to medical therapy, NOT standalone cure. Best outcomes: early AGA (Norwood II–IV; Ludwig I–II) + concurrent medical therapy (topical or oral minoxidil; finasteride 1 mg/day or dutasteride 0.5 mg/day in males; spironolactone in females). PRP on its own in advanced AGA (Norwood V–VII) produces minimal results. Set patient expectations accordingly. | Alves/Grimalt RCT 2018; Hausauer RCT 2018; combination therapy literature
GEOGRAPHY NOTE — Spain-specific context | In Spain, PRP is classified as a procedure using a medical device (centrifuge + tubes). It falls under the competence of licensed physicians (médicos). Requirements: CE-marked centrifuge and tubes; informed consent specifying off-label aesthetic use; clinical facility with minimum equipment for managing adverse events (anaphylaxis protocol, despite very low risk for autologous). The term 'PRP' is not controlled as a brand name — any practitioner may claim to offer PRP; kit quality and protocol matter. | AEMPS guidance; Spanish Real Decreto 1591/2009 (Medical Devices)
Fidia RichPlat: fidiaprp.com = US ORTHOPEDIC product (RichPlat Ortho). Fidia (Italy; also Hyalsystem fillers + Nithya PDRN) likely has an AESTHETIC RichPlat — verify Fidia Italy/AEMPS. Added via refresh-loop from user tip.
REGULATORY — AEMPS closed-system list (Spain) | See sheet '09 AEMPS Closed Systems'. AEMPS classifies PRP-obtaining device KITS via the CCPS application into 'técnica cerrada' (no environmental exposure of blood/product) and 'con exposición ambiental mínima' (exposure only at the initial anticoagulant-load and/or the final PRP/PPP-extraction steps). This is a Spain-specific MEDICAL-DEVICE classification (dynamic list; act. 17-feb-2025) — separate from CE marking and FDA. Being on the list means the kit can obtain PRP as a 'closed system'; it does NOT confer an aesthetic indication (PRP aesthetic use in Spain remains off-label with a CE-marked device + informed consent). | AEMPS CCPS list
CLASSIFICATION — three AEMPS-listed products are NOT platelet-rich plasma | Do not use these as aesthetic PRP: (1) Orthokine (Rhein/Orthogen, ref F0005) = autologous conditioned serum (ACS) — cell-free anti-inflammatory cytokines (IL-1Ra), glass-bead incubation, orthopedic. (2) QREM Cytokine (refs M01-001/M02-001) = autologous cytokine-rich serum (SARC), also cell-free — no platelet concentrate delivered. (3) WINT PRP (Biovico, ref PRO00579) = a DKK-1/Wnt-inhibition intra-articular OA product. All three appear on the AEMPS PRP-systems list but are mechanistically distinct from PRP. | Manufacturer + peer-review
INTRASPIN — conditional 'closed technique' status | IntraSpin (L-PRF, BioHorizons/Intra-Lock, ref ISS220) qualifies as 'técnica cerrada' per AEMPS ONLY if the L-PRF clot is administered directly; any additional manipulation reclassifies it as an OPEN system. It is an L-PRF membrane/clot device (glass tubes, no anticoagulant), not liquid PRP. | AEMPS + manufacturer
CORRECTION — Dr.PRP manufacturer + FDA (logged) | Dr.PRP corrected via AEMPS refresh-loop: legal manufacturer is Rmedica Co. Ltd. (NOT 'AHC' as previously catalogued; US distributor Dr. PRP USA LLC). Device is double-compartment dual-centrifugation (double-spin), not a single-spin gel-separator tube. FDA 510(k) now CONFIRMED cleared: BK230818 (20 cc, 2023) + BK251201 (30 cc). | FDA CBER + KMDA
OEM CLUSTER — T-Biyoteknoloji / T-LAB (Turkey) | KIT PRP T-LAB (T-1000), PRPHD (HD1000), Curevital (CU500), Next PRP (N500) and Dives Med are the SAME Turkish OEM (T-Biyoteknoloji; CE 2195/Szutest). Two protocol families: (a) gel-free buffy-coat tubes with selectable LP/LR PRP (T-1000, HD1000, Dives Med); (b) single-spin 1500g×4min pure-PRP (CU500, N500). No independent RCTs → Class IV, Low confidence. | Manufacturer IFUs (ifuport.com)
BTI ENDORET kit family = one PRGF platform | The six AEMPS-listed ENDORET kits (estética KMU10/-SCP/-GEL; dental KMU15/-PLUS/KMU-16; oftalmología KMU11/KMU18; traumatología EDK1/EDK2/EDK1-SCP; dermatología-TOPIC KMU10-TPC; KIT 19 KMU19) are ALL the single PRGF-Endoret platform (Na-citrate 3.8%, ~580g×8min, leukocyte-poor P-PRP, ~2–3× — correction from prior 1.6×). Only ESTÉTICA and DERMATOLOGIA/TOPIC are facial-aesthetic; dental/ophthal/trauma are not; KMU19 application unconfirmed. | BTI + Anitua peer-review

## AEMPS Closed Systems
AEMPS — Listado OFICIAL de productos sanitarios considerados sistemas de obtención de PRP CERRADOS (comunicados en la aplicación CCPS). Grupo de Trabajo de PRP de la AEMPS. Última actualización del listado: 17-feb-2025. LISTADO DINÁMICO — verificar la versión vigente en aemps.gob.es antes de cualquier uso regulatorio o clínico. 'Técnica cerrada' = obtención del PRP sin exposición ambiental del material de partida ni del producto. 'Con exposición ambiental mínima' = exposición sólo en etapas iniciales (p.ej. carga del anticoagulante) y/o finales (extracción del PRP/PPP); extremar precauciones (*). Notas al pie: (*) En las fases con posible exposición ambiental deben extremarse las precauciones y realizarse en entornos adecuados para reducir al mínimo el riesgo de contaminación. (**) Si la carga de anticoagulante se hace mediante técnica cerrada se considera SIN exposición ambiental; el uso de aguja y ampollas NO se considera técnica cerrada. (***) Los centros deben obtener el PRP según las IFU y disponer de los componentes NO incluidos en el kit para mantener las condiciones por las que se clasificaron. Fuente: https://www.aemps.gob.es/medicamentos-de-uso-humano/medicamentos-biologicos/plasma-rico-en-plaquetas-prp/listado-de-productos-sanitarios-considerados-como-sistemas-de-obtencion-de-prp-cerrados-comunicados-en-la-aplicacion-ccps/
Nombre comercial (AEMPS) | Referencias incluidas | Fabricante | País | Clasificación AEMPS | Observaciones / Momentos de posible exposición ambiental | Observaciones contenido kit | Equivalente en catálogo (01) | Notas
3E-PRP | 3E-10; 3E-20 | Pervice Korea Co., Ltd. | Rep. de Corea | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** | No incluye anticoagulante ni aguja para su carga | NUEVO — 3E-PRP
Accelerate | 620-12-02; 620-22-06; 620-22-08 | EmCyte Corporation | EEUU | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** | No incluye anticoagulante ni sistema para extracción de sangre | NUEVO — EmCyte Accelerate | Casa EmCyte (ya en catálogo con PurePRP / GenesisCS)
BS PRP SYSTEM | BS PRP-20 | BS Medical Co, Ltd. | Rep. de Corea | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** | No incluye anticoagulante ni sistema para extracción de sangre | NUEVO — BS PRP System
CUREVITAL PRP | CU500 | T-Biyoteknoloji Laboratuvar Estetik Medikal Kozmetik San. TIC. Ltd. | Turquía | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** y finales | No incluye agujas | NUEVO — Curevital PRP (casa T-Biyoteknoloji)
DISPRAS | P7-1212; P7-1222; P7-1232; P7-1242; P7-1212-SB; P7-1222-SB | Lab. Fidia Farmacéutica SLU | España | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** | No incluye el anticoagulante | NUEVO — Dispras (Fidia)
DIVES MED PRP COMPLETE SOLUTION | 4251871102881 | T-Biyoteknoloji Laboratuvar Estetik Medikal Kozmetik San. TIC. Ltd. | Turquía | Cerrado con exposición ambiental mínima | Etapas finales | Kit completo | NUEVO — Dives Med PRP (casa T-Biyoteknoloji)
Dr.PRP | DrPRP | Rmedica Co. Ltd. | Rep. de Corea | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** y finales | Kit completo | Dr.PRP | AEMPS: fabricante = Rmedica (catálogo actualizado desde 'AHC')
ENDORET KIT 19 | KMU19 | BTI Biotechnology Institute SL | España | Cerrado con exposición ambiental mínima | Etapas finales | Kit completo | PRGF-Endoret (BTI)
ENDORET KIT DERMATOLOGIA (TOPIC LINE) | KMU10-TPC | BTI Biotechnology Institute SL | España | Cerrado con exposición ambiental mínima | Etapas finales | Kit completo | PRGF-Endoret (BTI) — variante dermatología
ENDORET® KIT DENTAL DESECHABLE | KMU15; KMU15-PLUS; KMU-16 | BTI Biotechnology Institute SL | España | Cerrado con exposición ambiental mínima | Etapas finales | Kit completo | PRGF-Endoret (BTI) — variante dental
ENDORET® KIT OFTALMOLOGÍA | KMU11; KMU18 | BTI Biotechnology Institute SL | España | Cerrado con exposición ambiental mínima | Etapas finales | Kit completo | PRGF-Endoret (BTI) — variante oftalmología
Hy-Tissue PRP | P7-4050; P7-4020 | Lab. Fidia Farmacéutica SLU | España | Cerrado (técnica cerrada) | Solo si uso de jeringa precargada con anticoagulante (C-Lock 4% 5 mL) no incluida en kit, y que debe facilitar la empresa distribuidora. En caso contrario, se considera como con exposición ambiental mínima en fases iniciales. | No incluye el anticoagulante (C-Lock 4% 5 mL) | NUEVO — Hy-Tissue PRP
Hy-Tissue PRP | P7-4050; P7-4020 | Lab. Fidia Farmacéutica SLU | España | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante con técnica no cerrada, etc)**. | No incluye el anticoagulante | NUEVO — Hy-Tissue PRP
INTRASPIN | ISS220 | BioHorizons Implant Systems, Inc. | EEUU | Cerrado (técnica cerrada) | Solo si administración directa del coágulo obtenido. Si se dan operaciones adicionales se considerará un sistema abierto. | Kit completo | NUEVO — IntraSpin (L-PRF) | Dispositivo L-PRF (BioHorizons/Intra-Lock); base del protocolo IntraSpin/Fibrin
Jeringa Doble ACP | ABS-10014 | Arthrex Inc. | EEUU | Cerrado (técnica cerrada) | Si carga de anticoagulante mediante técnica no cerrada, se considerará con exposición ambiental mínima en fases iniciales. | No incluye ningún sistema para extracción de sangre | Arthrex ACP (Autologous Conditioned Plasma) | Mismo producto/referencia aparece también en el listado 'exposición ambiental mínima'
Jeringa Doble ACP | ABS-10014 | Arthrex Inc. | EEUU | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** | No incluye anticoagulante, aguja para su carga, ni sistema para extracción de sangre | Arthrex ACP (Autologous Conditioned Plasma)
KIT ENDORET® ESTÉTICA DESECHABLE | KMU10; KMU10-SCP; KMU10-GEL | BTI Biotechnology Institute SL | España | Cerrado con exposición ambiental mínima | Etapas finales | Kit completo | PRGF-Endoret (BTI) — variante estética | ENDORET = plataforma PRGF de BTI (mismo sistema; kits por aplicación)
KIT ENDORET® TRAUMATOLOGÍA | EDK1; EDK2; EDK1-SCP | BTI Biotechnology Institute SL | España | Cerrado con exposición ambiental mínima | Etapas finales | Kit completo | PRGF-Endoret (BTI) — variante traumatología
KIT PRP T-LAB | T-1000 | T-Biyoteknoloji Laboratuvar Estetik Medikal Kozmetik San. TIC. Ltd. | Turquía | Cerrado con exposición ambiental mínima | Etapas finales | Kit completo | NUEVO — T-LAB PRP (casa T-Biyoteknoloji) | OEM turco T-LAB — fabrica también DIVES MED / PRPHD / CUREVITAL / Next PRP
Next PRP | N500 (94223N500) | T-Biyoteknoloji Laboratuvar Estetik Medikal Kozmetik San. TIC. Ltd. | Turquía | Cerrado con exposición ambiental mínima | Etapas finales | Kit completo | NUEVO — Next PRP (casa T-Biyoteknoloji)
OLIN-1 | OL-1999; OL-1994 | Enelini Associats, SLU. | España | Cerrado (técnica cerrada) | Solo si uso de jeringa precargada con anticoagulante y conector para su transferencia a jeringa para extracción de sangre (componentes no incluidos en kit que debe facilitar la empresa distribuidora). En caso contrario, se considera como con exposición ambiental mínima en fases iniciales. | No incluye anticoagulante ni sistema de extracción (ver observación / listado 2) | NUEVO — OLIN-1
OLIN-1 | OL-1999; OL-1994 | Enelini Associats, SLU. | España | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante con técnica no cerrada, etc)**. | No incluye anticoagulante, aguja para su carga, ni sistema para extracción de sangre | NUEVO — OLIN-1
Orthokine | F0005 | Rhein Device GmbH | Alemania | Cerrado con exposición ambiental mínima | Etapas finales (si uso de agujas para extracción del PPP o PRP obtenido) | No incluye sistema para extracción de sangre ni productos para extraer el PPP y PRP obtenidos | NUEVO — Orthokine (ACS/Orthokine) | Sistema Orthokine/ACS (Rhein/Orthogen) — suero autólogo condicionado; uso mayoritario ortopédico
PRO PRP | P30A; P10; P60 | Goodmorning Bio CO., Ltd. | Rep. de Corea | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** y finales | No incluye el anticoagulante | NUEVO — Pro PRP
PROLO 30 PRP | A001-0101 | RM Life Science CO., Ltd. | Rep. de Corea | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** y finales | Kit completo | NUEVO — Prolo 30 PRP
Prosys ACT10 Bio Kit | Prosys ACT10 | Medisarang Co, Ltd. | Rep. de Corea | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** y finales | Kit completo | NUEVO — Prosys ACT10
PRPHD PRP KIT | HD1000 | T-Biyoteknoloji Laboratuvar Estetik Medikal Kozmetik San. TIC. Ltd. | Turquía | Cerrado con exposición ambiental mínima | Etapas finales | Kit completo | NUEVO — PRPHD (casa T-Biyoteknoloji)
QREM CYTOKINE | M01-001; M02-001 | Tecnología Regenerativa QREM SL | España | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** | No incluye anticoagulante, ni jeringas y conectores | NUEVO — QREM Cytokine
Regen ACR-C CLASSIC | R-ACR C1/B | Regen Lab SA | Suiza | Cerrado (técnica cerrada) |  | Kit completo | NUEVO — Regen ACR-C (línea ACR)
Regen ACR-C EXTRA | R-ACR C2/B | Regen Lab SA | Suiza | Cerrado (técnica cerrada) | Solo si utilización según fase 4a y 5 descrita en las IFUs. En caso contrario, se considera como con exposición ambiental mínima en fases finales. | Kit completo | NUEVO — Regen ACR-C (línea ACR) | Misma referencia aparece también en 'exposición ambiental mínima' (fase 4b)
Regen ACR-C EXTRA | R-ACR C2/B | Regen Lab SA | Suiza | Cerrado con exposición ambiental mínima | Etapas finales si utilización según fase 4b descrita en las IFUs | Kit completo | NUEVO — Regen ACR-C (línea ACR) | Misma referencia figura también como 'técnica cerrada' (fase 4a y 5)
Regen ACR-C PLUS | R-ACR C/BA | Regen Lab SA | Suiza | Cerrado con exposición ambiental mínima | Etapas finales | Kit completo | NUEVO — Regen ACR-C (línea ACR)
RegenKit-BCT | RK-BCT-1-2 | Regen Lab SA | Suiza | Cerrado (técnica cerrada) | Si aspiración del PPP obtenido (fase 4b de las IFU), se considerará con exposición ambiental mínima en fases finales. | Kit completo | RegenKit-BCT
RegenKit-BCT | RK-BCT-1-2 | Regen Lab SA | Suiza | Cerrado con exposición ambiental mínima | Etapas finales si utilización según fase 4b descrita en las IFUs | No incluye aguja larga para extracción del PPP | RegenKit-BCT
RegenKit-BCT | RK-BCT-T; RK-BCT-3 | Regen Lab SA | Suiza | Cerrado con exposición ambiental mínima | Etapas finales. Si no se usa un sistema cerrado para extracción de sangre, también en fases iniciales. | No incluye sistema para extracción de sangre ni productos para la extracción del PPP y PRP obtenido | RegenKit-BCT-3 / RegenKit-THT | RK-BCT-3 = variante multi-tubo; RK-BCT-T asociada a activador THT
REIN PRP | OM-RE1-0001 | Kaizen Endüstri ve Sanayi Dis Ticaret Ltd. | Turquía | Cerrado con exposición ambiental mínima | Etapas finales | Kit completo | NUEVO — Rein PRP
S&M PRP MONOCAP | STR-PRPKIT-01 | Str Biyo Teknolojileri Medical IML.ITH | Turquía | Cerrado con exposición ambiental mínima | Etapas finales | Kit completo | NUEVO — S&M PRP Monocap
SBL.PRP.3 15 ml | SBLxyzz | D-Biotech SRL | Italia | Cerrado (técnica cerrada) | No incluye vacutainer con aguja de mariposa, soporte de extracción por vacío y jeringa 10 mL L-L, componentes necesarios indicados en las IFU para obtener el PRP mediante técnica cerrada. Estos componentes los debe facilitar la empresa distribuidora. En caso contrario, se considerará como con exposición ambiental mínima en fases finales. | No incluye sistema de extracción de sangre (ver observación) | NUEVO — SBL.PRP.3 | 'SBLxyzz' figura así (aparente marcador de referencia genérica) en el listado AEMPS
SBL.PRP.3 15 ml | SBLxyzz | D-Biotech SRL | Italia | Cerrado con exposición ambiental mínima | Etapas finales (si uso de agujas para extracción del PPP o PRP obtenido) | No incluye sistema para extracción de sangre ni productos para extraer el PPP y PRP obtenidos | NUEVO — SBL.PRP.3
SW-PRP KIT CERRADO PARA PRP | SW-PRP | Seawon Meditech Co., Ltd. | Rep. de Corea | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** | No incluye anticoagulante, agujas ni jeringas | NUEVO — SW-PRP
TriCell Best | Tricell New (TEQAGRU0001) / Tricell New S (TEQAGRU0002) | REV-MED Inc. | Rep. de Corea | Cerrado con exposición ambiental mínima | Etapas iniciales (carga anticoagulante, etc)** y finales | No incluye la jeringa para extraer el PRP obtenido | NUEVO — TriCell Best
Tropocells PRP | 11MACD7SGB | Estar Technologies Ltd | Israel | Cerrado con exposición ambiental mínima | Etapas finales (siguiendo opciones cerradas de las IFU para recogida del PRP) | Kit completo | Tropocells PRP Kit
VIVOSTAT PRF | VS406; VS400; VS410; VS422 | Vivostat A/S | Dinamarca | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** y finales | Kit completo | NUEVO — Vivostat PRF | Sistema automatizado PRF autólogo (no PRP líquido clásico)
WINT PRP | PRO00579 | Biovico Sp.z.o.o | Polonia | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** y finales | No incluye el anticoagulante | NUEVO — WINT PRP (casa Biovico)
Xerthra PRP KIT | PRO00447 | Biovico Sp.z.o.o | Polonia | Cerrado con exposición ambiental mínima | Etapas iniciales (carga de anticoagulante, etc)** y finales | No incluye el anticoagulante | NUEVO — Xerthra (casa Biovico)

## Dose per Zone
PRP DOSING — volume PER POINT + total per session by indication. PRP is autologous so there is no 'dilution' (final PRP volume depends on the kit/blood draw — see sheet 01 'Final Volume'); what matters clinically is how much to place per point and the point spacing. Volumes are CONSENSUS ranges; deliver as a course (usually 3–4 sessions, monthly, then maintenance). Activate (CaCl2/thrombin) only if the protocol requires it. See sheet 02 (Indication × Product) for LP vs LR choice per zone.
Zone / Indication | Volume per point | Point spacing / number | Total per session | Depth / plane | Technique | Notes | Confidence
Acne scars | 0.05–0.1 mL | over scar bed + subcision points | 1–3 mL | Intradermal / subdermal (at scar base) | PRP + microneedling / subcision (combination = Class A) | Best combined with RF-microneedling or subcision | Consensus
Combination (PRP + HA / biostimulator) | per base product | per base product | PRP 1–2 mL added | per base product plane | Structural product first; PRP intradermal after (separate layer) | Do not premix LP-PRP with cross-linked HA | Consensus
Face — global rejuvenation / skin quality | 0.05–0.1 mL | microdroplet grid (nappage) | 2–4 mL full face | Intradermal / superficial subdermal | Nappage microdroplets or linear threads (cannula) | LP preferred (less bruising); combine with microneedling | Consensus
Neck / décolleté | 0.05–0.1 mL | microdroplet grid | 2–4 mL per area | Intradermal nappage | Microdroplet nappage | Skin-quality/crepiness; series of sessions | Consensus
Scalp — AGA (androgenetic alopecia) | 0.05–0.1 mL | ~1 cm grid (0.5–1 cm) | 3–6 mL per full scalp | Intradermal to subdermal (nappage or bolus) | 30 G needle papules or nappage across affected zone | Strongest-evidence PRP indication; LP or LR; monthly × 3–4 then q3–6 mo | Consensus (Class A for AGA)
Under-eye / infraorbital | 0.1–0.2 mL per side | 1–3 depots or fanning | 0.3–0.6 mL per side | Deep subdermal / supraperiosteal (superficial to orbicularis) | Slow retrograde depots or cannula; aspirate | LP-PRP; conservative; pair with HA for hollowing | Consensus